Instructions for Calibration of TSI Nephelometer using CPD and TSI software

The instructions will permit you to accomplish a neph calibration at SPO assuming the following are true:

- The TSI nephelometer is connected to port USB0 (Port 1 on USB-to-serial converter box)

- The CO2 delivery line is connected to the ¼” Swagelok bulkhead fitting on the nephelometer comms panel.

A calibration typically takes 30-40 minutes to complete.

Instructions

- With a cpdclient session running (blue window with aerosol data visible), single-click the terminal icon to open up a terminal window.

- Type ‘dosneph’ (type what is inside the quotes).  Executing this script does several things.  It changes the configuration of the cpd software so that it will not try to communicate with the neph anymore.  It stops and restarts CPD with this new configuration.  The blue cpdclient window disappears (although CPD is still running in the background).  It opens the TSI nephelometer software through a DOS emulator.

- Single-click on the TSI software window to make it the active window.  At the Main Menu in the TSI Software window, choose ‘com port and data path’.  If the COM port is not already set to COM2, please do that and press <Tab> to get to ‘OK’, then press <Enter> (the ‘<>’ is a designation of a key to be pressed).  The com port setting should already be on COM2 – this is the default.  Note:  You may have a different version of the software than I have to test here.  You may have to put the comm port/data path to COM1.  You should just choose the one that actually communicates with the neph.  If you can’t get it to work, we will stop there and regroup.

- At the Main Menu, choose ‘Calibration’.  This will take you to the calibration screen.  At the top of the Calibration screen, the current values of the K2 and K4 constants are shown.  You will be determining the new values of these constants through the calibration.  Now you will go down the list of steps in the TSI neph software window to complete the calibration.  If at any time you make a mistake, do something out of order, etc., you can press <F10> to exit out of the calibration and start again.

- Press <1>.  This will show you a list of the gases you may use for the low span gas.  By default, ‘Air’ has been chosen.  If ‘Air’ is not highlighted, please choose ‘Air’.  Then press <Enter>.  A check mark will appear next to Item #1 in the list to show that this step has been completed.

- Press <2>.  This will show you a list of the gases you may use for the high span gas.  By default, ‘CO2’ has been chosen.  If ‘CO2’ is not highlighted, please choose ‘CO2’.  Then press <Enter>.  A check mark will appear next to Item #2 in the list to show that this step has been completed.

- Press <3>.  This will change the position of the filter valve to the filtered air position.  A check mark will appear next to Item #3 in the list to show that this step has been completed.

- Press <4>.  A check mark will appear next to Item #4 in the list.  This starts data collection on filtered air.  Eventually, you will see the low span gas numbers in the window start changing.  Allow to run on Step #4 for a minimum of 300 seconds (a timer is visible in the window) to purge the instrument with filtered air.

- Press <5>.  A check mark will appear next to Item #5 in the list.  This clears the data buffer and restarts data collection on filtered air.  The timer reverts back to ‘0’ seconds.  Allow to run on Step #5 for a minimum of 300 seconds.  This is the actual calibration measurement on filtered air.

- Press <6>.  A check mark will appear next to Item #6 in the list.  This will store the filtered air measurement data.  The low span gas numbers stop changing when this is done.

- Press <7>.  A check mark will appear next to Item #7 in the list.  This command does nothing if you do not have the neph blower installed.  If you do, it toggles the blower ‘on’ or ‘off’.

- Press <8>.  A check mark will appear next to Item #8 in the list.  Since we are in a different configuration than that described in the TSI Nephelometer Manual, we do not need to manually plug the neph inlet and outlet.  You will have to manually seal off the vacuum from the neph outlet however by closing the rotameter (downstream of the neph outlet) needle valve.  If you do not do this, the pump will suck the CO2 out of the neph faster than it will be fed in in the following steps.

- Press <9>.  A check mark will appear next to Item #9 in the list.  This is to verify that you are feeding CO2 into the nephelometer.  This will be accomplished using the following steps:

Open the main valve on the CO2 tank.  CO2 is now being fed into the neph.  This would normally be exhausted to the pump, but since you closed the large rotameter valve it will be exhausted out of the neph inlet.  Make sure the flow rate of CO2 is 5-6 lpm.

- Press <A>.  A check mark will appear next to Item A in the list.  This starts data collection on filtered CO2.  Eventually, you will see the high span gas numbers in the window start changing (sometimes takes up to 30 sec).  Allow to run on Step A for a minimum of 300 seconds to purge the instrument with filtered air.

- Press <B>.  A check mark will appear next to Item B in the list.  This clears the data buffer and restarts data collection on filtered CO2.  The timer reverts back to ‘0’ seconds.  Allow to run on Step B for a minimum of 300 seconds.  This is the actual calibration measurement on filtered CO2.

- Press <C>.  A check mark will appear next to Item C in the list.  This will store the filtered CO2 measurement data.  The high span gas numbers stop changing when this is done.  At this point, new K2 and K4 constants will appear at the top of the screen.  If these change by a large amount (>5-10%), you may wish to repeat the calibration (you should email me and we can discuss if the cal looks acceptable).  A span check should be performed after every calibration to spot any potential calibration problems (more on this later).  You can now close the main valve on the CO2 tank.

- Press <D>.  This allows you to enter a label that will be stored with this calibration that will allow users to see when, where, and by whom this nephelometer was last calibrated.  We recommend the following protocol:

Date         Neph Serial #      who@location        high + low span gases

An example of this label would be:

100728a   #1022    JFB@SPO   CO2+Air

The ‘a’ on the date entry signifies it was the first calibration performed that day.  Successive calibrations on that date would be identified as 100728b, 100728c, etc.  The initials are JFB and you performed this calibration at ‘SPO’.  When you have finished writing the label, press <Enter> to store it.  A check mark will appear next to Item D in the list.

- If the calibration was acceptable, press <E>.  This will save the new constants to the nephelometer.  Please write down the current (i.e., old) and new K2 and K4 constants for each channel and send them to me in an email.

- Single-click on the TSI software window.  Press <F> to initiate a filtered air background check after opening the rotameter needle valve to the normal flow.  This will take ~ 7 more minutes (60 seconds for a purge, 300 seconds for the measurement, 60 seconds for a second purge).  When this is complete a check mark will appear next to Item F on the list.  You can confirm that the background measurement is completed when the green ‘valve’ light on the nephelometer (next to the serial port connector) is illuminated.

- Press <F10> key to move out of the calibration window.  Answer ‘Y’ to “Exit to Menu?”.

- At main menu, choose ‘Exit’ and the TSI software window will close.  Single-click on the icon “Stop CPD”, followed by choosing “restart” to bring CPD back.

- After I verify that the new constants look acceptable, you should perform a span check.  Hopefully the span check will have small errors for each channel and a small average error.  If we actually manage to make things worse, we can always put the old PMT back in and manually re-enter the old calibration settings.