Aerosol System Daily Operations Records

MONTH: 					YEAR:						STATION: CMDL/MAUNA LOA OBSERVATORY

| DAY | System vacuum gauge ~10* | 4 rotometers ~30 lpm | CN drier flow ~7 lpm* | CNC alcohol level OK # | Change PSAP filter if <0.80 | Change PSAP reference filter (Once/week) | Neph span check (Once/2 weeks) | Impactor Film change (Once/2 weeks | Comments |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |

* Pump may be down, contact CMDL scientist.				If you have any problems call:

# Please fill CN Counter with Butyl Alcohol				Anne Jefferson	303-497-6493

+ Please change PSAP filter.							Betsy Andrews	303-497-5171

## Typical Parameter Values at MLO

| Window | <Alt> key | Parameter | Typical | Lower limit | Upper limit | Action |
| --- | --- | --- | --- | --- | --- | --- |
| Aerosol Sampling Status | A | RH_sample | 40 | 0 | 45 | Check inlet heater |
|  | A | T_sample | 25 | 10 | 40 |  |
|  | A | T_rack | 30 | 20 | 35 | Check air conditioner |
|  | A | T_pumpbox | 32 | 25 | 40 |  |
|  | A | Q_analyzer | 30 | 28 | 32 | Check pump, adjust flow |
| Aerosol dP sensors | A | dP_Neph_imp | 2 – 10 | 2 | 13 | Check for leaks around impactors |
|  | A | dP_CNC_vac | 400 | 350 | 500 | Check CN pump |
|  | A | dP_pitot | 70 | 60 | 80 | Check carbon vane pump |
| CN counter | A | Drier flow | 8 | 7.5 | 8.5 | Adjust valve |
| Best UPS | B | Inverter | OFF |  |  |  |
|  |  | Charger | ON |  |  |  |
|  |  | AC Volts IN | 113 | 110 | 120 | Fix generators |
|  |  | AC Volts OUT | 121 | 118 | 125 |  |
|  |  | AC Amps OUT |  |  |  |  |
|  |  | Output Load VA |  |  |  | Check that all systems are drawing power, and than nothing unauthorized is connected to UPS |
|  |  | Battery DC V | 13.9 | 13 | 14 | Possible UPS malfunction |
|  |  | Frequency | 60 | 50 | 70 | Fix generators |
| PSAP | L | Flow | 3.0 | 2.5 | 3.5 | Adjust flow |
|  | L | Transmittance | 0.9 | 0.7 | 1.0 | Change filter |
|  | L | I_sample | 250000 | 90000 | 300000 | Check light source |
|  | L | I_reference | 570000 | 250000 | 300000 | Check light source |
|  | L | Filter change | FALSE | FALSE | FALSE | You forgot to press <alt><L><M><3> after last filter change. |

| Window | <Alt> | Parameter | Typical | Lower | Upper | Action |
| --- | --- | --- | --- | --- | --- | --- |
| Neph_N | N | Bkg_tot_Blue | 2.0E-6 | 1E-7 | 5*typ. |  |
|  |  | Bkg_tot_Green | 3.0E-6 | 1E-7 | 5*typ. |  |
|  |  | Bkg_tot_Red | 5.0E-6 | 1E-7 | 5*typ. |  |
|  |  | Bkg_back_Blu | 4.0E-6 | 1E-7 | 5*typ. |  |
|  |  | Bkg_back_Grn | 5.0E-6 | 1E-7 | 5*typ. |  |
|  |  | Bkg_bak_Red | 8.0E-6 | 1E-7 | 5*typ. |  |
|  |  | MODE | NBXX | BBXX | ZBXX | Any other value is an error |
|  |  | P_hPa | 680 | 660 | 700 |  |
|  |  | SecToGo | 30000 |  |  | Must decrease by one every second |
|  |  | Sample_T | 300 | 290 | 320 | Check room temperature |
|  |  | Inlet_T | 297 | 290 | 320 | Check room temperature |
|  |  | Next_Z | Xx:57:00 |  |  | Possible error in CP.INI entry for [Neph] ZeroAt= entry. |
|  |  | GreenRef | 4.0E+5 | 3.0E+5 | 5.0E+5 | Lamp may be failing. |
|  |  | Lamp_V | 13 | 10 | 14 | Lamp may be failing |
|  |  | Lamp_I | 5.7 | 5 | 7.5 | Lamp_V times Lamp_I should be 75.0 |
| uMAC | U | Room Temp | 38 | 30 | 40 | Check room temperature |
|  |  | Power supply | 5.0 | 4.9 | 5.1 | Power supply bad? |

The 7-wavelength Aethalometer according to Betsy  ---- DRAFT!!!

Specifications

Model Number	AE-8 (actually it’s an AE-31 Multi-wavelength instrument)

Serial Number		890501

CMDL Number	CD0000323164

Security Code		111 (need this to change certain parameters)

Table 1 Channel assignments for each wavelength and associated sigmas

| ***** | channel | 1 | 2 | 3 | 4 | 5 | 6 | 7 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| NEW | Wavelength (nm) | 370 | 470 | 521 | 565 | 700 | 430 | 950 |
|  | Sigma (m2/g) | 39.5 | 31.1 | 28.1 | 25.9 | 20.9 | 34.0 | 15.4 |
| ***** | *********** | **** | **** | **** | **** | **** | **** | **** |
| OLD | Wavelength (nm) | 370 | 470 | 521 | 590 | 660 | 880 | 950 |
|  | Sigma (m2/g) | 39.5 | 31.1 | 28.1 | 24.8 | 22.2 | 16.6 | 15.4 |

Area of filter spot: 0.5189 cm2 (manufacturer value: 0.5 cm2)

Flow range: 0-10 slpm. Note: flow must be greater than 1 lpm for measurements.

Default max attenuation before filter change: 75 (BC loading of 4 g/cm2)

Measurement noise level: 1 ng/cm2 which is equivalent to ~50 ng/m3 (at 5 min freq.)

Sampling frequency (Timebase): 2 min - 1 h

Operation

- Instrument takes ~30 minutes to warm up

- To stop quickly – switch off (keeps data, but doesn’t write summary file)

- To stop normally – press <STOP>, observe screen, press <STOP>, enter security code (111<enter>). You may need to hold the <STOP> button down as a simple double click doesn’t seem to get you to the screen where you can enter the security code.

- Data disk can be exchanged at anytime without interfering with measurements.

- It appears the only time that the aethalometer can be stopped gracefully is when it is in measurement mode.  If it freaks out because of high BC concentrations or if it’s in the middle of some sort of internal test you either need to cycle the power or (in the case of the test) wait until the test is over. (GEESH!)

- The tape advances every time you begin normal sampling mode (e.g., after quitting normal sampling mode by method 2 or 3).  I haven’t done an automatic tape advance.  You can also set the tape advance time to be at a specific time period.

Main System Menu

The up and down arrows are used to scroll through the Main System Menu.  The enter key is used to choose a menu option.  You have several options in the Main System Menu.  Anytime you start up the instrument or leave normal sampling mode you go to the Main System Menu.  The initial option is Operate.  You can use the up and down arrows to get to the other menu options or you can hit <enter> to choose Operate and return to sampling mode.

Operate - Automatic Mode versus Manual Mode

When the instrument first starts up, the instrument asks whether you wish to go into ‘automatic mode’. If you don’t do anything, after about 1 minute the instrument goes into automatic mode by default and starts making measurements in ~15 min after doing some internal checks.

If you hit enter (which selects Operate from the Main System Menu), you again get asked if you wish to go into ‘automatic mode’.  If you use the down arrow to select NO, then you go into ‘manual mode’. Manual mode means that you have to hit a couple buttons (follow the prompts on the screen) before the instruments starts its internal checks and then begins sampling.

Change Settings Menu

This menu lets you (surprise) change some of the operating parameters.  See manual for extensive details. Below are listed what I think are probably reasonable choices.

Table 2 Change Settings Menu and suggested values

| Setting | Choice | Range and comments |
| --- | --- | --- |
| Time and Date |  |  |
| Timebase (sample frequency) | 2 min | 2-60 min |
| Tapesaver | 0 – off | 0,1,3 (Off, x3, x10) |
| Analog output port/Alarm | 1 – alarm | 0-analog, 1-alarm |
| Analog out channel | 1 | 1 thru 7 (doesn’t matter for us) |
| Warm up wait | 0 – No | 0, 1 (0-avoids long warm up) |
| COM mode | Dataline | Dataline, Gesytec, off |
| COM settings | 9600,8,1,None | Various |
| Overwrite old data | 0 – No | 0,1 (0-no overwrite) |
| Filter change time | 0 – disable | I would change filter for each test |
| Security Code | 111 | Don’t change! |
| Data format | 0 – US | 0,1 (US, Euro) |
| BC display | 0 – ng/m3 | 0,1 (ng/m3, ug/m3) |
| Data format | 0 – extended | 0,1 (extended, compressed) |
| UV channel | 0 – UV on | 0,1 (UV on, UV off) |
| Instrument serial # | 890 | Not sure why it said 999 when we got |
| Gesytic id | ?? | Not applicable to us |
| Instrument type | 2 – 7 wavel. | 0,1,2 (Ae-1x,Ae-2x,Ae-3x) |
| Sigma for lamps | see Table 1 | Different for old and new light source |
| Spots per advance | 1 | 1,2 (spots closer, spots farther apart) |

Notes:

- The values for the settings in Table 2 are given in the ae_setup.txt file written to disk (they aren’t logged by the log_port program).

- There is also a message file (mf*.txt) written to disk that gives information about the run which may or may not be useful. See manual on page 82 for more info on both these files; examples of both the ae_setup.txt and mf*txt files are also attached to this document.

- The wavelengths of the lamps are hardwired into the aethalometer output files.  The wavelength is not used in the calculation of BC, only the sigma’s are, but don’t be confused if you look at raw aethalometer files and the stated wavelengths are for the old LED source.

Signals and flow

Choosing this takes you to a screen that displays sensor and reference signals for one lamp and the flow. (More info on p. 64.)  You can also test the bypass valve, but we aren’t using it so why bother! (

Self tests

This option takes about 5 minutes and runs through a series of internal tests.  It’s not an automatic procedure because the user is required to respond to prompts on the screen occasionally.  This is one of those processes which can’t be interrupted while it is occurring except by cycling the power.  More information is in the manual on page 65.

-lamps		-analog out		-COM port

-pumps	-screen/front LEDs	-tape advance

Calibrate flowmeter

See manual for description of procedure

Flow volume units

I have it set at standard units right now.  We need to correct their standard temperature of 20 C to our standard temperature of 0 C.

Software upgrade

Optical Test

Install new tape

See manual p. 68-72.

Data file format

The aethalometer manual has the WRONG data file format description for its bc*.csv data output files!!

The correct expanded data file format is (for the new wavelength source):

Date,time,bc_370,bc_470,bc_521,bc_565,bc_700,bc_430,bc_950, flow,

sz_370,sb_370,rz_370,rb_370,bypass fraction, cumulative_dATN_370,

sz_470,sb_470,rz_470,rb_470,bypass fraction, cumulative_dATN_470,

sz_521,sb_521,rz_521,rb_521,bypass fraction, cumulative_dATN_521,

sz_565,sb_565,rz_565,rb_565,bypass fraction, cumulative_dATN_565,

sz_700,sb_700,rz_700,rb_700,bypass fraction, cumulative_dATN_700,

sz_430,sb_430,rz_430,rb_430,bypass fraction, cumulative_dATN_430,

sz_950,sb_950,rz_950,rb_950,bypass fraction, cumulative_dATN_950

The time in the data file format corresponds to the START of the measurement cycle.

It is unclear what the ‘cumulative_dATN_#’ field really contains, but it looks like some number that increases as the filter loading increases.

Screen display (p. 84 of manual)

Line 1 - system time and date

Line 2 – tape remaining and whether tape saver is on

Line 3 – disk space remaining (like how much longer the tests can run) and flow (lpm)

Line 4 – value of BC in ng/m3 (or g/m3) – this line is an AVERAGE of the measurements from all 7 channels.

Page 84 also describes the panel LEDs

Calculations

Attenuation: ATN = -100 ln((sb-sz)/(rb-rz))

sb = sensor beam signal, sz = sensor zero signal

rb = reference beam signal, rz = reference zero signal

typical zero signal values are ~0.02

typical beam signal values are in range 1-2

Delta Attenuation: dATN = ATN(t1)-ATN(t0)

Pseudo-Absorption (m-1): p_bap = dATN*(Area/(Flow*Time))*(1/1000)			     Area is filter area in cm2, Flow is in lpm, Time is in minutes. (factor of (1/1000) gets units to work, takes into account 100 from ATN calculation and 1/10 from converting flow and area units)

BC concentration (ng m-3): [bc] = dATN*(Area/(Flow*Time*Sigma))*1e6	     Area is filter area in cm2, Flow is in lpm, Time is in minutes, Sigma is in m2/g. (factor 1e6 gets units to work, takes into account 100 from ATN calculation, 1/10 from converting flow and area units and converting from g to ng)

CMDL processing

CMDL takes the expanded aethalometer output file described above and uses /aer/prg/perl/cnvtka_ to make 3 file types (ka_, kb_ and kc_) in standard CMDL format.

ka_ file format is:

STN,Year,Start_UTC,Flag,blank,blank,ATN370,p_bap370,ATN470,p_bap470,

ATN521,p_bap521,ATN565,p_bap565,ATN700,p_bap700,ATN430,p_bap430,ATN950,p_bap950,flow,area,cum_dATN_565,sigma1,sigma2,sigma3,sigma4,sigma5,

sigma6,sigma7,ID,sd_STN,sd_Year,sd_Start_UTC,sd_Flag,sd_blank,sd_blank,sd_ATN370,sd_p_bap370,sd_ATN470,sd_p_bap470,sd_ATN521,sd_p_bap521,

sd_ATN565,sd_p_bap565,sd_ATN700,sd_p_bap700,sd_ATN430,sd_p_bap430,

sd_ATN950,sd_p_bap950,sd_flow,sd_area,sd_cum_dATN_565,sd_sigma1,

sd_sigma2,sd_sigma3,sd_sigma4,sd_sigma5,sd_sigma6,sd_sigma7

The numbers 1-7 correspond to the seven wavelength channels. ATN_#, dATN_# and p_bap are calculated as described above. The flow is in lpm, the area is in cm2. The cum_delta_attenuation_4 is what I think is the cumulative change in attenuation for the 565 nm wavelength channel. The sigma values are those given for the new light source.

kb_ file format is:

STN,Year,Start_UTC,Flag,blank,blank,BC370,BC470,BC521,BC565,BC700,BC430,BC950, sd_STN,sd_Year,sd_Start_UTC,sd_Flag,sd_blank,sd_blank, sd_BC370,sd_BC470,sd_BC521,sd_BC565,sd_BC700,sd_BC430,sd_BC950

kc_ file format is:

STN,Year,startDOY,Flags,Q_slpm,,Bap370,Bap470,Bap520,Bap564,Bap699,Bap430,Bap949,ATN370,ATN470,ATN520,ATN564,ATN699,ATN430,ATN949,segID,STN,Year,startDOY,sd_Flags,sd_Q_slpm,,sd_Bap370,sd_Bap470,sd_Bap520,sd_Bap564,sd_Bap699,sd_Bap430,sd_Bap949,sd_ATN370,sd_ATN470,sd_ATN520,sd_ATN564,sd_ATN699,sd_ATN430,sd_ATN949,

XSHOW

The ka_ and kb files are ingested into xshow and the data are plotted on the Aeth_Status plot.  There are 3 panes:

- the first pane shows the aethalometer flow in lpm

- the second plot shows the Pseudo_absorption values in Mm-1 for the 430, 565, and 700 nm channels

- the third plot shows the cumulative_delta_attenuation_4 value to give an indication of when the filter was changed.

# Text boxes

Tony Hansen is the man in charge of aethalometer manufacturing and is a good person to contact for trouble shooting if all else fails.

Tony Hansen’s email and phone number are:

TonyHansen@mageesci.com

(510) 845-2801

The webpage http://www.mageesci.com/ for the aethalometer also provides info and useful downloads