DAILY CHECKLIST – CMDL Aerosol System

Month:	________Year:200__												Station: Trinidad Head, California

| Date | CNC butanol level: fill y/n | CN drier flow* ~8.5 lpm | CNC vacuum 14 ” Hg | CNC status lights green | PSAP filter change | PSAP flow ~1.2 lpm | Q_stack* ~850 lpm | Water reservoir level: fill y/n | MFC box, hum. rack ~29 lpm | MFC box, filt. rack ~29 lpm | Filter vacuum ” Hg | Comments (e.g., rain, smoke, instrument issues, vehicle traffic, etc.etc.) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 |  |  |  |  |  |  |  |  |  |  |  |  |
| 2 |  |  |  |  |  |  |  |  |  |  |  |  |
| 3 |  |  |  |  |  |  |  |  |  |  |  |  |
| 4 |  |  |  |  |  |  |  |  |  |  |  |  |
| 5 |  |  |  |  |  |  |  |  |  |  |  |  |
| 6 |  |  |  |  |  |  |  |  |  |  |  |  |
| 7 |  |  |  |  |  |  |  |  |  |  |  |  |
| 8 |  |  |  |  |  |  |  |  |  |  |  |  |
| 9 |  |  |  |  |  |  |  |  |  |  |  |  |
| 10 |  |  |  |  |  |  |  |  |  |  |  |  |
| 11 |  |  |  |  |  |  |  |  |  |  |  |  |
| 12 |  |  |  |  |  |  |  |  |  |  |  |  |
| 13 |  |  |  |  |  |  |  |  |  |  |  |  |
| 14 |  |  |  |  |  |  |  |  |  |  |  |  |
| 15 |  |  |  |  |  |  |  |  |  |  |  |  |
| 16 |  |  |  |  |  |  |  |  |  |  |  |  |
| 17 |  |  |  |  |  |  |  |  |  |  |  |  |
| 18 |  |  |  |  |  |  |  |  |  |  |  |  |
| 19 |  |  |  |  |  |  |  |  |  |  |  |  |
| 20 |  |  |  |  |  |  |  |  |  |  |  |  |
| 21 |  |  |  |  |  |  |  |  |  |  |  |  |
| 22 |  |  |  |  |  |  |  |  |  |  |  |  |
| 23 |  |  |  |  |  |  |  |  |  |  |  |  |
| 24 |  |  |  |  |  |  |  |  |  |  |  |  |
| 25 |  |  |  |  |  |  |  |  |  |  |  |  |
| 26 |  |  |  |  |  |  |  |  |  |  |  |  |
| 27 |  |  |  |  |  |  |  |  |  |  |  |  |
| 28 |  |  |  |  |  |  |  |  |  |  |  |  |
| 29 |  |  |  |  |  |  |  |  |  |  |  |  |
| 30 |  |  |  |  |  |  |  |  |  |  |  |  |
| 31 |  |  |  |  |  |  |  |  |  |  |  |  |

Problems: Contact Pat Sheridan (303) 497-6672 or Betsy Andrews (303) 497-5171			Logging Status: ABCFHLMNOUVW

*Check values on computer screen in ‘aerosol window’ <alt><a>

Month:________________  Year: 200___

WEEKLY TASKS

| PSAP ref. filter Date Time (UTC) |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Neph span check Date Time (UTC) |  |  |  |  |  |
| Impactor cleaning (Hum. Rack) Date |  |  |  |  |  |
| Carousel filters (Filter Rack) Date |  |  |  |  |  |
| Leak check Date | ambCN= amb_dP= filtCN= filt_dP= | ambCN= amb_dP= filtCN= filt_dP= | ambCN= amb_dP= filtCN= filt_dP= | ambCN= amb_dP= filtCN= filt_dP= | ambCN= amb_dP= filtCN= filt_dP= |
| Mail carousel filters to PMEL |  |  |  |  |  |
| Pumpbox Vac guage~0.25”H2O Rotameters ~30 lpm |  |  |  |  |  |

MONTHLY TASKS

| Mail PSAP filters, checklists to NOAA |  |
| --- | --- |
| Flush CN Butanol |  |
| Check supply levels | Butanol DI water PSAP filters Impactor films Plastic bags Other:___________________ |

Typical Parameter Values

| Window | <Alt> key | Parameter | Typical | Lower limit | Upper limit | Action |
| --- | --- | --- | --- | --- | --- | --- |
| Aerosol Sampling Status | A | RH_sample | 40 | 0 | 45 | Check inlet heater |
|  | A | T_sample | 25 | 10 | 40 |  |
|  | A | T_rack | 30 | 20 | 35 | Check air conditioner |
|  | A | T_pumpbox | 32 | 25 | 40 |  |
|  | A | Q_analyzer | 30 | 28 | 32 | Check pump, adjust flow |
| Aerosol dP sensors | A | dP_Neph_imp | 2 – 10 | 2 | 13 | Check for leaks around impactors |
|  | A | dP_CNC_vac | 700 | 650 | 750 | Check CN pump |
|  | A | dP_pitot | 70 | 60 | 80 | Check carbon vane pump |
| CN counter | A | Drier flow | 8 | 7.5 | 8.5 | Adjust valve |
| Best UPS | B | Inverter | OFF |  |  |  |
|  |  | Charger | ON |  |  |  |
|  |  | AC Volts IN | 113 | 110 | 120 | Fix generators |
|  |  | AC Volts OUT | 121 | 118 | 125 |  |
|  |  | AC Amps OUT |  |  |  |  |
|  |  | Output Load VA |  |  |  | Check that all systems are drawing power, and than nothing unauthorized is connected to UPS |
|  |  | Battery DC V | 13.9 | 13 | 14 | Possible UPS malfunction |
|  |  | Frequency | 60 | 50 | 70 | Fix generators |
| PSAP | L | Flow | 3.0 | 2.5 | 3.5 | Adjust flow |
|  | L | Transmittance | 0.9 | 0.7 | 1.0 | Change filter |
|  | L | I_sample | 250000 | 90000 | 300000 | Check light source |
|  | L | I_reference | 570000 | 250000 | 300000 | Check light source |
|  | L | Filter change | FALSE | FALSE | FALSE | You forgot to press <Esc>ILF after last filter change. |

| Window | <Alt> | Parameter | Typical | Lower | Upper | Action |
| --- | --- | --- | --- | --- | --- | --- |
| Neph_N & Neph_O | N & O | Bkg_tot_Blue | 2.0E-5 | 1E-7 | 5*typ. |  |
|  |  | Bkg_tot_Green | 1.8E-5 | 1E-7 | 5*typ. |  |
|  |  | Bkg_tot_Red | 1.6E-5 | 1E-7 | 5*typ. |  |
|  |  | Bkg_back_Blu | 6.0E-6 | 1E-7 | 5*typ. |  |
|  |  | Bkg_back_Grn | 5.0E-6 | 1E-7 | 5*typ. |  |
|  |  | Bkg_bak_Red | 4.0E-6 | 1E-7 | 5*typ. |  |
|  |  | MODE | NBXX | BBXX | ZBXX | Any other value is an error |
|  |  | P_hPa | 1000 | 950 | 1020 |  |
|  |  | SecToGo | 30000 |  |  | Must decrease by one every second |
|  |  | Sample_T | 300 | 290 | 320 | Check room temperature |
|  |  | Inlet_T | 297 | 290 | 320 | Check room temperature |
|  |  | Next_Z | Xx:57:00 |  |  | Possible error in CP.INI entry for [Neph] ZeroAt= entry. |
|  |  | GreenRef | 4.0E+5 | 3.0E+5 | 5.0E+5 | Lamp may be failing. |
|  |  | Lamp_V | 13 | 10 | 14 | Lamp may be failing |
|  |  | Lamp_I | 5.7 | 5 | 7.5 | Lamp_V times Lamp_I should be 75.0 |
| Humidograph | H | refNeph | 30% | 0% | 40% |  |
|  |  | wetNeph | Varies | 0% | 95% |  |
|  |  | RH_S1 (RH after preheater) | 30% | 0% | 40% | Check humidograph system |
|  |  | T_S1 (T after preheater) | 297 | 290 | 320 | Check humidograph system |
|  |  | RH_S2 (RH after humidifier) | varies | 0% | 97% | Check humidograph system |
|  |  | T_S2 (T after humidifier) | 297 | 290 | 320 | Check humidograph system |
| uMAC & vMAC | U & V | Room Temp | 38 | 30 | 40 | Check room temperature |
|  |  | Power supply | 5.0 | 4.9 | 5.1 | Power supply bad? |

System Schematic

Humidifier system

RH sensor locations