3760 CN butanol drain connector replacement

- Let CN counter go dry by not filling it. (put note in the log that you are going to let it run dry).  You will know it’s dry when counts are zero and you can’t see any liquid at all in the window (even if you tilt the CN counter slightly).

- Pull CN counter out of rack.  To do this you’ll need to disconnect (a) power cord,  (b) sample inlet air, (c) vacuum, (d) electrical connection to the readout box with the red number and (e) butanol fill line.  (you’ll need a 9/16” wrench for the vacuum line and sample inlet line, you might also need a 1/2” wrench for the sample inlet line).  I’m guessing you’ll have to be behind the rack disconnect everything except the sample line.

- Wrap a piece of Teflon tape around the threads of the new style connector.

- Unscrew the drain connector (5/8” wrench) and replace with new one.  Note:  the drain connector you have might be an old style connector or a new style connector.  The old style connector looks similar to the new style connector except that it has a 3/8” extension on the side that screws into the CN counter.

- Before you reinstall CN counter in rack, add a little bit of butanol and make sure nothing leaks.

- Once the CN counter is re-installed put a note in the log saying the repair is done.