## Other Maintenance
Item | Description | Frequency
Inspect system | Replace broken tubing and hoses, make sure fittings are tight. | periodically, every month will suffice
Clean stack heater | Pull a clean damp cloth through tube to clean dust and grime deposits | once per year
Replace C vanes in pump | remove end plate and replace broken C vanes | as necessary
Replace filters in neph, impactor box | replace old filters with new ones from consumable parts inventory | as necessary
system leak check | put leak check filter on 2-inch stainless heater tube in stack, confirm that CN counts drop to low values (if CN counter samples from same inlet tube). | periodically, every month will suffice

## Calibrations
Calibration | Description | Required Tools | Documentation | Recommendation
PSAP flow calibration | calibration of flow rate for PSAP light absorption measurement | Flow calibrator, such as BIOS Definer or Gilibrator | http://www.biosint.com/Products/products2.php, http://www.sisweb.com/sptd/gilibrat.htm | at least once per year
Impactor box mass flow controller calibration | calibration of flow rate for impactor box MFC flow rate | Flow calibrator, such as BIOS Definer or Gilibrator | http://www.biosint.com/Products/products2.php, http://www.sisweb.com/sptd/gilibrat.htm | at least once per year
TSI neph calibration | calibration of TSI neph optics/detectors using CO2 gas | CO2 gas, TSI neph software | TSI Nephelometer Operating Manual, http://www.esrl.noaa.gov/gmd/aero/docs/index.html | as necessary
Differential Pressure sensor calibration | calibration of Neph Impactor dP sensor on umac box | Differential pressure gauge (e.g., magnehelic gauge, Model 2010, 0-10 in. H2O) | http://www.dwyer-inst.cn | at least once per year
Differential Pressure sensor calibration | calibration of System Vacuum dP sensor on umac box | Differential pressure gauge (e.g., magnehelic gauge, Model 2210, 0-10 psid) | http://www.dwyer-inst.cn | at least once per year
Pitot tube calibration | calibration of pitot tube to determine stack flow rate | Differential pressure gauge (e.g., magnehelic gauge, Model 2000-0, 0-0.5 in. H2O) | http://www.dwyer-inst.cn | at least once per year
Vaisala T/RH probe calibration | Calibration of sample line temperature and RH measurements | Trusted T and RH standards (e.g., calibrated thermometer, calibrated RH standard, salt solutions, etc. |  | at least once per year

## Spare Parts
Item | Vendor | Catalog # or Part # | Telephone# | Web Page | Recent Price | Recommended number to keep on hand | Last Known Price | Details
Temp/RH sensor w/ no cable | Vaisala China Ltd. | HMP60A12A0A1B0 | +86 10 5827 4100 | http://www.vaisala.com | 300 | 2 | 600
Carbon vane pump | Gast Manufacturing, Inc. (China distributor) | 0823-V103-G608X | +86-21-5241 5599 | http://www.gastmfg.com/asia.html | 700 | 1 | 700 | Can be configured for 120VAC or 240VAC operation
Pump repair kit | Gast Manufacturing, Inc. (China distributor) | K479A | +86-21-5241 5599 | http://www.gastmfg.com/asia.html | 65 | 2 | 130
Centrifugal blower | Ametek Beijing Rep. Office | 117636-01 | +86.10.8526.2111 ext. 13 | http://www.ametek.com.cn | 660 | 1 | 660 | 240VAC Model
Solenoid valve | ASCO Asia - Beijing | 8210G034 | +86 10 8563 1122 | http://www.ascovalve.net | 200 | 1 | 200 | Impactor box solenoid
Solenoid valve | ASCO Asia - Beijing | 8210G094 | +86 10 8563 1122 | http://www.ascovalve.net | 125 | 1 | 125 | Impactor box solenoid
16-port USB to serial box | VSCOM | USB-16COM-RM | sales@vscom.de | http://www.vscom.de/640.htm | 400 | 1 | 400 | Edgeport box replacement
Electronic ball valve | Swagelok Beijing Rep. Office | SS-63TS12-42ACE | +86-10-8203 0815 | http://www.swagelok.com/ | 700 | 1 | 700 | 220VAC, 50 Hz model, for impactor box
Brooks Mass Flow Controller | Brooks Instrument China | #5851EA14BW2E2BA with #320B203BMA process connectors | +86 21 3895 4788 | http://brooksinstrument.com | 2100 | 0/1 | $0/$2100 | Replacement for MFC in Impactor box

## Consumables
Item | Vendor | Catalog # or Part # | Telephone# | Web Page | Recent Price | Number to keep on hand | Cost | Details
PSAP Filters, 250/pk | VWR International | 21431-350 |  | https://www.vwrsp.com | 48 | 2 | 96 | For PSAP absorption measurement
CO2 gas | misc. vendors | Instrument Grade, 99.99% |  |  | 200 | 1 | 200 | For calibrating nephelometer
Silicone Spray | 3M, or get from hardware store | 3M I.D.# 62-4678-4930-3 |  | Go to 3M website, type in Product I.D. number | 5 | 2 | 10 | For spraying impaction surfaces in impactors
Dow Corning High Vacuum Grease | Fisher Scientific |  |  | http://www.fishersci.com | 30 | 1 | 30 | For o-ring seals
Q-tips | Many vendors |  |  |  | 3 | 1 | 3 | For cleaning surfaces
Gelman HEPA filter | VWR Scientific | Gelman Part# 12144, VWR Part# 28145-145 |  | http://www.vwrsp.com | 62 | 2 | 124 | Use to replace stack leak check filter, zero filter in neph, or filter in impactor box
Parker filter | Webster Associates, Inc. | IDN-4G X 10 (10 filter pack) | 800-627-4875 | http://www.filternetwork.com | 8.98 | 2 | 17.96 | To replace small filter in nephelometer
Tubing (includes stainless steel, conductive silicone black tubing, tygon, and polyethylene vacuum tubing) | misc. vendors |  |  |  | 300 | 1 | 300 | Tubing will degrade and crack over time. You need to keep a supply of vacuum and sample line tubing on hand to replace damaged tubing.
Ethanol/Methanol, bottle | misc. vendors |  |  |  | 20 | 1 | 20 | General cleaning