DAILY CHECKLIST – CMDL Aerosol System

Month:	________Year:200__													Station: Cape San Juan, Puerto Rico

| Date | CNC butanol level: fill y/n | CN drier flow* ~8 lpm | CNC vacuum >14 ” Hg | CNC status lights G/R/Both | PSAP filter change | PSAP flow ~0.5 lpm | Q_stack* ~850 lpm | MFC box, Neph. ~29 lpm | MFC box, filters ~29 lpm | Active filter (1-8) | Active filter pressure (60-100 hPa) | Comments (e.g., rain, smoke, instrument issues, vehicle traffic, etc.etc.) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 |  |  |  |  |  |  |  |  |  |  |  |  |
| 2 |  |  |  |  |  |  |  |  |  |  |  |  |
| 3 |  |  |  |  |  |  |  |  |  |  |  |  |
| 4 |  |  |  |  |  |  |  |  |  |  |  |  |
| 5 |  |  |  |  |  |  |  |  |  |  |  |  |
| 6 |  |  |  |  |  |  |  |  |  |  |  |  |
| 7 |  |  |  |  |  |  |  |  |  |  |  |  |
| 8 |  |  |  |  |  |  |  |  |  |  |  |  |
| 9 |  |  |  |  |  |  |  |  |  |  |  |  |
| 10 |  |  |  |  |  |  |  |  |  |  |  |  |
| 11 |  |  |  |  |  |  |  |  |  |  |  |  |
| 12 |  |  |  |  |  |  |  |  |  |  |  |  |
| 13 |  |  |  |  |  |  |  |  |  |  |  |  |
| 14 |  |  |  |  |  |  |  |  |  |  |  |  |
| 15 |  |  |  |  |  |  |  |  |  |  |  |  |
| 16 |  |  |  |  |  |  |  |  |  |  |  |  |
| 17 |  |  |  |  |  |  |  |  |  |  |  |  |
| 18 |  |  |  |  |  |  |  |  |  |  |  |  |
| 19 |  |  |  |  |  |  |  |  |  |  |  |  |
| 20 |  |  |  |  |  |  |  |  |  |  |  |  |
| 21 |  |  |  |  |  |  |  |  |  |  |  |  |
| 22 |  |  |  |  |  |  |  |  |  |  |  |  |
| 23 |  |  |  |  |  |  |  |  |  |  |  |  |
| 24 |  |  |  |  |  |  |  |  |  |  |  |  |
| 25 |  |  |  |  |  |  |  |  |  |  |  |  |
| 26 |  |  |  |  |  |  |  |  |  |  |  |  |
| 27 |  |  |  |  |  |  |  |  |  |  |  |  |
| 28 |  |  |  |  |  |  |  |  |  |  |  |  |
| 29 |  |  |  |  |  |  |  |  |  |  |  |  |
| 30 |  |  |  |  |  |  |  |  |  |  |  |  |
| 31 |  |  |  |  |  |  |  |  |  |  |  |  |

Problems: Contact Pat Sheridan (303) 497-6672 or Betsy Andrews (303) 497-5171			Logging Status: *ABCEGMNU

Boxes in grey are not currently applicable to the CSJ station.                                *Check values on computer screen in ‘aerosol window’ <alt+A>

Month:________________  Year: 200___

WEEKLY TASKS

| Neph span check Date Time (UTC) |  |  |  |  |
| --- | --- | --- | --- | --- |
| Impactor films Date |  |  |  |  |
| Leak check Date |  |  |  |  |

MONTHLY TASKS

| Mail checklists to NOAA |  |
| --- | --- |
| Check supply levels | Butanol Impactor films CO2 Other:___________________ |

Schematic of CSJ Aerosol Flow System

NOAA/CMDL Impactor Box Flow Schematic – Normal Operation

B = Manual Ball Valve (open position)

S/NO = Electronic Solenoid Valve, Normally Open (without power)

S/NC = Electronic Solenoid Valve, Normally Closed (without power)

NOAA/CMDL Impactor Box Flow Schematic – Impactor Bypass Operation (during impactor changes)

B = Manual Ball Valve (closed position)

S/NO = Electronic Solenoid Valve, Normally Open (without power)

S/NC = Electronic Solenoid Valve, Normally Closed (without power)

NOAA/CMDL Impactor Box Flow Schematic – Zero Check Operation (same as normal operation except zero check port is opened and connected to CPC using conductive tubing.

B = Manual Ball Valve (open position)

S/NO = Electronic Solenoid Valve, Normally Open (without power)

S/NC = Electronic Solenoid Valve, Normally Closed (without power)

NOAA/CMDL Impactor Box Flow Schematic – Span Check Operation

B = Manual Ball Valve (open position)

S/NO = Electronic Solenoid Valve, Normally Open (without power)

S/NC = Electronic Solenoid Valve, Normally Closed (without power)

Aerosol Impactor Box – Wiring Diagram

P1 = Amp, MateNLok, 2x3 shells w/ sockets (connector to electronic ball valve)

P2 & P3 = Amp, MateNLok, 1x3 shell w/ pins (connectors to solenoids)

P4 = MateNLok 3x3 shell w/ pins (front panel of impactor box)

(G) = Ground

(H) = Hot (=L, Line on some connectors)

(N) = Neutral

Wire color and gauge is expressed like G-18 (green, 18 AWG)

Switched A/C power – connections to front panel connector

Vaisala Panel Mount Pin-out

Pin	Parameter

1	RH

2	Temp

3	7-28 VDC

4	common

Vaisala Humitter Sensor  Pin-out

Wire	Parameter

BRN	RH

V	Temp

Y	7-28 VDC

G	common

## Typical Parameter Values at CSJ

| Window | <Alt> key | Parameter | Typical | Lower limit | Upper limit | Action |
| --- | --- | --- | --- | --- | --- | --- |
| Aerosol Sampling Status | A | RH_sample | 40 | 0 | 45 | Check inlet heater |
|  | A | T_sample | 25 | 10 | 40 |  |
|  | A | T_rack | 30 | 20 | 35 | Check air conditioner |
|  | A | T_pumpbox | 32 | 25 | 40 |  |
|  | A | Q_analyzer | 30 | 28 | 32 | Check pump, adjust flow |
| Aerosol dP sensors | A | dP_Neph_imp | 2 – 10 | 2 | 15 | Check for leaks around impactors |
|  | A | dP_CNC_vac | 500 | 350 | 800 | Check CN pump |
|  | A | dP_pitot | 70 | 60 | 80 | Check carbon vane pump |
| CN counter | A | Drier flow | 8 | 7.5 | 8.5 | Adjust valve |
| Best UPS | B | Inverter | OFF |  |  |  |
|  |  | Charger | ON |  |  |  |
|  |  | AC Volts IN | 113 | 110 | 120 | Fix generators |
|  |  | AC Volts OUT | 121 | 118 | 125 |  |
|  |  | AC Amps OUT |  |  |  |  |
|  |  | Output Load VA |  |  |  | Check that all systems are drawing power |
|  |  | Battery DC V | 13.9 | 13 | 14 | Possible UPS malfunction |
|  |  | Frequency | 60 | 50 | 70 | Fix generators |
| PSAP | L | Flow | 3.0 | 2.5 | 3.5 | Adjust flow |
|  | L | Transmittance | 0.9 | 0.7 | 1.0 | Change filter |
|  | L | I_sample | 250000 | 90000 | 300000 | Check light source |
|  | L | I_reference | 570000 | 250000 | 300000 | Check light source |
|  | L | Filter change | FALSE | FALSE | FALSE | You forgot to press <alt><L><M><3> after last filter change. |

| Window | <Alt> | Parameter | Typical | Lower | Upper | Action |
| --- | --- | --- | --- | --- | --- | --- |
| Neph_N | N | Bkg_tot_Blue | 2.0E-6 | 1E-7 | 5*typ. |  |
|  |  | Bkg_tot_Green | 3.0E-6 | 1E-7 | 5*typ. |  |
|  |  | Bkg_tot_Red | 5.0E-6 | 1E-7 | 5*typ. |  |
|  |  | Bkg_back_Blu | 4.0E-6 | 1E-7 | 5*typ. |  |
|  |  | Bkg_back_Grn | 5.0E-6 | 1E-7 | 5*typ. |  |
|  |  | Bkg_bak_Red | 8.0E-6 | 1E-7 | 5*typ. |  |
|  |  | MODE | NBXX | BBXX | ZBXX | Any other value is an error |
|  |  | P_hPa | 1000 | 960 | 1040 |  |
|  |  | SecToGo | 30000 |  |  | Must decrease by one every second |
|  |  | Sample_T | 300 | 290 | 320 | Check room temperature |
|  |  | Inlet_T | 297 | 290 | 320 | Check room temperature |
|  |  | Next_Z | Xx:57:00 |  |  | Possible error in CP.INI entry for [Neph] ZeroAt= entry. |
|  |  | GreenRef | 2.0E+5 | 1.0E+5 | 5.0E+5 | Lamp may be failing. |
|  |  | Lamp_V | 13 | 12 | 14 | Lamp may be failing |
|  |  | Lamp_I | 6.0 | 5.5 | 7.0 | Lamp_V times Lamp_I should be ~ 75 |
| uMAC | U | Room Temp | 38 | 30 | 40 | Check room temperature |
|  |  | Power supply | 5.0 | 4.8 | 5.2 | Power supply bad? |

# Text boxes

Splitter

(150 lpm flow)

T/RH sensor

Rotameter

Blower

Control Line (30 lpm)

Pump for excess flow

30 lpm

30 lpm

30 lpm

Impactor Box and Aerosol Rack System

(30 lpm)

Carbon vane pump

Stack bypass flow (850-1000 lpm)

Ball Valve

exhaust

exhaust

exhaust

Pitot tube

Impactor dP port

B

10 micron

Impactor

1 micron

Impactor

Electronic

Ball Valve

(closed = 1 um path, open = 10 um path)

Flow outlet to MFC and pump

CO2 entry/Zero check port

Aerosol outlet to nephelometer

Port to PSAP inlet

HEPA Filter

S/

NO

S/

NC

Aerosol sample inlet

Return flow from nephelometer

Impactor dP port

B

10 micron

Impactor

1 micron

Impactor

Electronic

Ball Valve

(closed = 1 um path, open = 10 um path)

Flow outlet to MFC and pump

CO2 entry/Zero check port

Aerosol outlet to nephelometer

Port to PSAP inlet

HEPA Filter

S/

NO

S/

NC

Aerosol sample inlet

Return flow from nephelometer

Impactor dP port

B

10 micron

Impactor

1 micron

Impactor

Electronic

Ball Valve

(closed = 1 um path, open = 10 um path)

Flow outlet to MFC and pump

CO2 entry/Zero check port

Aerosol outlet to nephelometer

Port to PSAP inlet

HEPA Filter

S/

NO

S/

NC

Aerosol sample inlet

Return flow from nephelometer

Impactor dP port

B

10 micron

Impactor

1 micron

Impactor

Electronic

Ball Valve

(closed = 1 um path, open = 10 um path)

Flow outlet to MFC and pump

CO2 entry (5 lpm)/Zero check port

Aerosol outlet to nephelometer

Port to PSAP inlet

HEPA Filter

S/

NO

S/

NC

Aerosol sample inlet

Return flow from nephelometer

P4-1

P2-1

P4-2

P4-3

P2-2

P2-3

N.O. Valve

P3-1

P3-2

P3-3

N.C.

Valve

(G) G-18

(H) B-18

(N) W-18

IEC

B-18

W-18

G-18

(L)

(N)

(G)

P4-5

P4-6

N/C

BRW-18

R-18

P1-1

P1-2

P1-3

P1-4

P1-5

P1-6

N

H

G

To switched a/c for solenoid valves

G

To switched a/c for ball valve

Top of instrument rack showing aethalometer, and CPC flow control box

Bottom of instrument rack showing mac (micromac) box, PID control box, and impactor box

Top view of instrument rack showing USB extension cable (white cable) and CO2 inlet for span checks.