## Slide 1: NOAA CLAP: Continuous Light Absorption Photometer
Why bother with a new instrument?
  lost data due to overloaded filters
  unheated sample block
  leak-prone filter holder
  "black box" internal processing of PSAP
User interface – KISS (Keep It Simple, Stupid!)
  filter change button and status indicator
  flow control
  primary user interface is through 'cpd' software
Data output optimized for machine processing
Firmware can be upgraded in the field, remotely from Boulder
Characterization
  internal construction (photos)
  noise vs. averaging time
  LED intensity
  planned open source release of drawings, schematics, code

## Slide 2: NOAA CLAP – Assembled Product

## Slide 3: NOAA CLAP – Light Source Internal View

## Slide 4: NOAA CLAP – Assembly Details

## Slide 5
Chart: LED spectral output - CLAP s/n 10.002

## Slide 6
Chart: CLAP and PSAP noise vs. Averaging Interval Smoothers: PSAP=4-sec boxcar, CLAP=3-sec 1-pole/4-pass
| 4.0 | 8.0 | 12.0 | 16.0 | 24.0 | 40.0 | 64.0 | 100.0 | 160.0 | 252.0 | 400.0 | 632.0 | 1000.0 | 1584.0 | 2512.0 | 3980.0
green PSAP#094 (0.11 @ 60s) | 1.186 | 0.6886 | 0.4691 | 0.3584 | 0.2482 | 0.1594 | 0.1108 | 0.08411 | 0.06464 | 0.05336 | 0.04705 | 0.04232 | 0.03996 | 0.03654 | 0.03598 | 0.03217
green PSAP#106 (0.09 @ 60s) | 0.9554 | 0.5436 | 0.3772 | 0.2936 | 0.2049 | 0.1324 | 0.09083 | 0.06542 | 0.05075 | 0.0404 | 0.0363 | 0.03231 | 0.03076 | 0.02712 | 0.0273 | 0.02427
blue CLAP#001 (0.18 @ 60s) | 1.764 | 1.621 | 1.478 | 1.224 | 0.8737 | 0.5851 | 0.3839 | 0.2552 | 0.1765 | 0.1188 | 0.06861 | 0.04877 | 0.0278 | 0.02213 | 0.01524 | 0.007857 | 0.008103 | 0.006172
green CLAP#001 (0.15 @ 60s) | 1.737 | 1.59 | 1.441 | 1.187 | 0.8278 | 0.5572 | 0.3521 | 0.2336 | 0.151 | 0.09032 | 0.06628 | 0.04183 | 0.02521 | 0.01685 | 0.01081 | 0.01083 | 0.007141 | 0.007368
red CLAP#001 (0.11 @ 60s) | 1.343 | 1.224 | 1.115 | 0.906 | 0.6303 | 0.4156 | 0.2657 | 0.1704 | 0.112 | 0.07656 | 0.05218 | 0.03017 | 0.01851 | 0.01326 | 0.008158 | 0.006151 | 0.01106 | 0.004577

## Slide 7: CLAP vs. PSAP Comparison
sap (Mm-1)
CLAP (new)
PSAP (old)