## Slide 1: Aerosol Monitoring Workshop
Betsy Andrews, John Ogren, Patrick Sheridan
NOAA/ESRL/GMD
May 17-18, 2012

## Slide 2: Logistical Issues
Wireless network access uses same authentication as in GMAC meeting room
  Network: dsrc
  User/Password: gmac/gmac
Location of rest rooms
Please don’t wander around the building without an escort
Presentations can be downloaded from ftp://ftp.cmdl.noaa.gov/aerosol/etc/workshop_2012

## Slide 3: Participants
Station | Names
ALT | Sangeeta Sharma. Dan Veber
APP | James Sherman, Brett Taubman, Yong Zhou
CPR | Ian Gutierrez Molina
LLN | George (Neng-Huei) Lin, Eric (Hao-Ping) Chia
MLO | John Barnes
GMD | Betsy Andrews, Patrick Sheridan. John Ogren

## Slide 4: Schedule
Date Time Room/Leader | Topic
5/17 13:30-15:00 GB514 / JAO | General discussion: Science : objectives, strategies, instruments Operations: common computer tasks, supplies, spares, consumables
5/17 15:00-15:30 | break
5/17 15:30-17:00 3D401 / PJS | Nephelometer, part 1
5/18 08:30-10:00 3D401 / PJS | Nephelometer, part 2
5/18 10:00-10:30 | break
5/18 10:30-12:00 GB514 / EJA | Data analysis (data extraction, averaging, exporting; using templates for statistical analyses and plots, documentation and FAQs)
5/18 12:00-13:00 | lunch
5/18 13:00-14:30 3D401 / PJS | CPC
5/18 14:30-15:00 | break
5/18 15:00-16:30 GB514 / JAO | CLAP/PSAP (flow calibration, spot area measurement, understanding flags, troubleshooting)
5/18 16:30-1700 GB514 / JAO | Wrap-up discussion

## Slide 5: NOAA-federated Long-term Aerosol Network
http://www.esrl.noaa.gov/gmd/aero/
Legend NOAA affiliate future sites
ALT
SPO
SMO
MLO
BRW
BND
SGP
THD
CPT
WLG
CPR
TIK
SUM
LLN
KPS
SPL
WHI
ETL
EGB
AMY
ARN
APP
BEO
GSN
Notes: Future sites are ones that we expect to integrate into the NOAA-affiliated network within the next 5 years
We welcome more collaborations, especially in Central and South America, Africa, Central and South Asia

## Slide 6: Growth of NOAA Aerosol Network
Chart: year
| 1973.0 | 1974.0 | 1975.0 | 1976.0 | 1977.0 | 1978.0 | 1979.0 | 1980.0 | 1981.0 | 1982.0 | 1983.0 | 1984.0 | 1985.0 | 1986.0 | 1987.0 | 1988.0 | 1989.0 | 1990.0 | 1991.0 | 1992.0 | 1993.0 | 1994.0 | 1995.0 | 1996.0 | 1997.0 | 1998.0 | 1999.0 | 2000.0 | 2001.0 | 2002.0 | 2003.0 | 2004.0 | 2005.0 | 2006.0 | 2007.0 | 2008.0 | 2009.0 | 2010.0 | 2011.0 | 2012.0 | 2013.0
Stations | 0 | 2 | 2 | 3 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 5 | 5 | 7 | 7 | 7 | 7 | 7 | 8 | 7 | 8 | 8 | 8 | 10 | 13 | 15 | 14 | 17 | 20 | 21 | 24 | 24 | 0

## Slide 7: DAY 1: OPERATIONS DISCUSSION
common computer tasks
  LiveCPD
  AER_VM
  Help systems, FAQ
  Interpreting new data report
  Annual clean up of lost data
Annual maintenance procedures
  Questions?
supplies, spares, consumables
  Any problems?
Funding issues
  What can we do to support each other?
    Acquire/maintain funds
    Conserve funds

## Slide 8: Remote Access to LiveCPD
Text-mode
  ssh [-p port#] cpd@IP.add.re.ss
Graphical mode
  svnc [-p port#] cpd@IP.add.re.ss
Default port# is 22 (ssh), port forwarding on firewall might require a different port#
Firewall must be configured to allow remote access to LiveCPD system
Use of tunneling system
  lstunnels
  connect_tunnel2 STN

## Slide 9: Remote Access to LiveCPD
AER_VM has software for text-mode access
  should we add graphical mode software?

## Slide 10: Sending data directly to AER_VM server
Advantages
  restricted data access (no data are sent to NOAA)
  local data processing during field campaigns
  'sync' script only synchronizes programs
  all data processing is done on AER_VM system
Disadvantages – no NOAA support for
  troubleshooting data problems
  quicklook plots
Requirements
  server must run continuously
  LiveCPD must have direct path to server
    both on same local network, or
    server has a public IP address and firewall allows direct access on port 22

## Slide 11: LiveCPD -> AER_VM Data Flow
Site operator (local or remote)
Field system
NOAA app server
WDCA archive
ssh +
VNC
World Data Center for Aerosols
NOAA servers
sync
AER_VM cpx2
data.*
LiveCPD cpd
nfs
ftp
rsync
“aerolab6”
scp
programs only

## Slide 12: WMO Rolling Review of Requirements
Observation requirements
  http://www.wmo.int/pages/prog/www/OSY/GOS-RRR.html
  Observing requirements database, http://www.wmo-sat.info/db/
  WMO and GCOS will be updating the requirements database, based on community input
Review of capabilities
Statement of Guidance
Repeat the above three steps periodically

## Slide 13: Observation Requirements Database
Table below is the result of a search for “Aerosol and Radiation” requirements in “Lower Troposphere”
Who wants to contribute to fixing this?

## Slide 14: DAY 1: GENERAL DISCUSSION
Establishing new stations
  In undersampled regions
Is there desire to have a repository for sharing overheads for presentations?
  E.g., network map, instrument descriptions, summary descriptive statistics
Any other topics to discuss?

## Slide 15: DAY 1: SCIENCE DISCUSSION
What “non-climate” users are interested in our data?
  Chemical weather forecasting
  Sand and dust storm warnings
  How do we meet those needs?
Strategies for addressing science questions
  Multi-station analyses (transport, transformation, trends, contrasts)
  GAW trends paper (21 stations, plus IMPROVE)
  Uncertainties (e.g., question from IPCC, “what is the uncertainty of an observational-based estimate of direct aerosol radiative forcing?”

## Slide 16: DAY 1: SCIENCE DISCUSSION
Instruments
  What are the weaknesses of our systems?
    CPC response falloff, butanol contamination
    Neph truncation correction
    PSAP/CLAP corrections
      Sensitivity to particle size
  What should we add next?
    Photoacoustic absorption?
    Extinction?
    Size distribution?
    Thermo-denuder? (e.g., COSMOS)

## Slide 17: Direct Aerosol Forcing of Climate
What are we after?
  Change in surface and top-of-atmosphere energy budget due to scattering and absorption of sunlight by anthropogenic aerosols, for the entire globe.
What do we need to measure?
  AEROSOL AMOUNT
    Anthropogenic aerosol optical depth, a
    Chemical measurements to determine anthropogenic fraction
  AEROSOL PROPERTIES
    absorption vs. scattering
    scattering vs. angle
    water uptake
What is our strategy?
  Models+satellites provide global distribution of AMOUNT
  Monitoring from surface and airplanes gives
    validation data for models and satellites
    climatology of aerosol PROPERTIES
  Radiative transfer models derive forcing from AMOUNT and PROPERTIES
Radiative Forcing Efficiency

## Slide 18: Variations in Aerosol Amount and Type
Extinction (Mm-1)
Single scattering albedo
AMOUNT
TYPE
A rich data set for evaluating chemical transport models
1 Mm-1 = 1 mm2 m-3, extinction cross-section per cubic meter of air
Percentiles 95 75 50 25 5
log scale!
Notes: Explain SSA
Single-scattering albedo used as a simple indicator of aerosol type
Emphasize variability at a given station, as well as among stations
Note how collatorative stations (yellow) greatly expand the comparison data

## Slide 19: NOAA CLAP: Continuous Light Absorption Photometer
Why bother with a new instrument?
  lost data due to overloaded filters
  unheated sample block
  leak-prone filter holder
  "black box" internal processing of PSAP
User interface – KISS (Keep It Simple, Stupid!)
  filter change button and status indicator
  flow control
  primary user interface is through 'cpd' software
Data output optimized for machine processing
Firmware can be upgraded in the field, remotely from Boulder
Characterization
  internal construction (photos)
  noise vs. averaging time
  LED intensity
  planned open source release of drawings, schematics, code

## Slide 20: NOAA CLAP – Assembled Product

## Slide 21: NOAA CLAP – Light Source Internal View

## Slide 22: NOAA CLAP – Assembly Details

## Slide 23: CLAP Deployment Plans (27 inst.)
Currently deployed at 17 stations
  ALT, AMF, AMY, APP, ARN, BND, BRW, CPR, CPT, ETL, GSN, LLN, MLO, SGP, SPL, SUM, THD
Planned deployments (6 stations)
  BEO, EGB, KPS (June, 2012)
  WHI (later in 2012)
  Two new Canadian stations (2013?)
Hot spare, field campaign (2 instruments)
With thermo-denuder (GSN, TBD)
  Compare with COSMOS and unheated CLAP

## Slide 24: CLAP Deployment
May, 2012
Legend NOAA affiliate future sites CLAP
ALT
SPO
SMO
MLO
BRW
BND
SGP
THD
CPT
WLG
CPR
TIK
SUM
LLN
KPS
SPL
WHI
ETL
EGB
AMY
ARN
APP
BEO
GSN
Notes: Future sites are ones that we expect to integrate into the NOAA-affiliated network within the next 5 years
We welcome more collaborations, especially in Central and South America, Africa, Central and South Asia

## Slide 25: Measurement of CLAP spot size
Automated analysis of test filters from each CLAP
Measurement uncertainty is 2%
Average area = 19.9 ± 0.5 mm2 (n=200)

## Slide 26: Measured LED spectral output
Chart: wavelength (nm)
CLAP 10.002
Wavelength (nm) | Blue | Green | Red
Peak | 461.4 | 522.2 | 653.5
Centroid (n=25) | 466.6 ±1.2 | 528.6 ±0.9 | 653.1 ±0.9
FWHM | 26 | 40 | 20

## Slide 27: CLAP noise (Mm-1) vs. averaging time
1
0.001
0.1
0.01
Boxplots represent 80 spots. Symbols show average noise of 6 PSAPs.

## Slide 28: CLAP noise vs. averaging time
1
0.001
0.1
0.01
Boxplots represent 194 CLAP detectors. Symbols show average noise of 6 PSAPs.
averaging time
standard deviation (Mm-1)

## Slide 29: CLAP precision analysis
Boxplots show variability of 7 instruments operated in parallel.
light absorption coefficient (Mm-1)
standard deviation (Mm-1)
0.6
0.4
0.2
0.0
0
5
10
15

## Slide 30: CLAP vs. PSAP Comparison at BND
sap (550 nm) (Mm-1)
CLAP (new)
PSAP (old)

## Slide 31: CLAP Flags
Flag bit | set by inst? | set by cpd | Description
0x8000 | no | no | reserved for future correction flag
0x4000 | no | no | CTS correction applied
0x2000 | no | no | Bond correction applied
0x1000 | no | no | Weiss correction applied
0x0800 | no | no | Not used by CLAP
0x0400 | yes | no | Case temperature unstable
0x0200 | yes | no | Temperature error (inlet or block)
0x0100 | yes | yes | Lamp or filter error
0x0080 | no | yes | Transmittance IrR_cAxx < 0.5
0x0040 | no | yes | Transmittance IrR_cAxx < 0.7
0x0020 | no | yes | Transmittance IrG_cAxx < 0.5
0x0010 | no | yes | Transmittance IrG_cAxx < 0.7
0x0008 | no | yes | Transmittance IrB_cAxx < 0.5
0x0004 | no | yes | Transmittance IrB_cAxx < 0.7
0x0002 | yes | no | Flow error
0x0001 | yes | yes | Filter changing

## Slide 32: CLAP Troubleshooting
Case temperature unstable
  Isolate case from the cool laboratory
  Change setpoint (35-40ºC)
Flowrate unstable
  Check pump
Unexplained filter changes
  Contact Derek
Poorly-defined spots
  Check proper sealing
  Make sure top half is properly oriented

## Slide 33: CLAP Flowmeter Calibration
Choosing the setpoint (Vstn = 1.0 lpm)
  Setpoint in slpm will be less than 1.0 at most sites
  Vstp = Vstn * (Pstn/Pstp) * (Tstp/Tstn)
Obtaining calibration data
  Step flowrate in 0.1 lpm increments from 0.2 to 1.5 lpm
  Record voltage and flow calibrator reading (adjusted to STP, 1013.25 hPa and 0ºC)

## Slide 34: CLAP Flowmeter Calibration
Calculating calibration correction
  Full-range calibration curve (3rd order) Instruments;A11;!Cal;QHardware,-0.37582;0.56804;-0.25566;0.0595
  Fine-tuning multiplier at normal setpoint Instruments;A11;!Cal;Q,0.988
Chart: CLAP flow calibration
| 1.01 | 1.11 | 1.2 | 1.310999999999997 | 1.395999999999997 | 1.494 | 1.603 | 1.702 | 1.802999999999997 | 1.903999999999995 | 2.001 | 2.105 | 2.202 | 2.299 | 2.398 | 2.498 | 2.599 | 2.703000000000001 | 2.804 | 2.898 | 2.996999999999998 | 3.1 | 3.194 | 3.297 | 3.398999999999999 | 3.5 | 3.597 | 3.705 | 3.799 | 3.9
Q_slpm | 0 | 0.02 | 0.041 | 0.064 | 0.082 | 0.102 | 0.124 | 0.144 | 0.165 | 0.187 | 0.208 | 0.239 | 0.266 | 0.296 | 0.333 | 0.376 | 0.426 | 0.478 | 0.53 | 0.58 | 0.636 | 0.702 | 0.768 | 0.844 | 0.928 | 1.02 | 1.12 | 1.244 | 1.359 | 1.492

## Slide 35: SUPPLEMENTARY MATERIAL

## Slide 36: GCOS 10 Monitoring Principles
1. The impact of new systems or changes to existing systems should be assessed prior to implementation.
2. A suitable period of overlap for new and old observing systems is required.
3. The details and history of local conditions, instruments, operating procedures, data processing algorithms and other factors pertinent to interpreting data (i.e., metadata) should be documented and treated with the same care as the data themselves.

## Slide 37: GCOS 10 Monitoring Principles
4. The quality and homogeneity of data should be regularly assessed as a part of routine operations.
5. Consideration of the needs for environmental and climate-monitoring products and assessments, such as IPCC assessments, should be integrated into national, regional and global observing priorities.
6. Operation of historically-uninterrupted stations and observing systems should be maintained.

## Slide 38: GCOS 10 Monitoring Principles
7. High priority for additional observations should be focused on data-poor regions, poorly observed parameters, regions sensitive to change, and key measurements with inadequate temporal resolution.
8. Long-term requirements, including appropriate sampling frequencies, should be specified to network designers, operators and instrument engineers at the outset of system design and implementation.

## Slide 39: GCOS 10 Monitoring Principles
9. The conversion of research observing systems to long-term operations in a carefully-planned manner should be promoted.
10. Data management systems that facilitate access, use and interpretation of data and products should be included as essential elements of climate monitoring systems.
http://www.wmo.ch/pages/prog/gcos/index.php?name=ClimateMonitoringPrinciples