## summaryO3 18 Jun 16
Averages of total ozone:
With the exisiting Calibration, Lamp tests included
Average is over the complete Mu range
D083_TO3_C | D083_TO3_D | D083_TO3_A | D083_TO3_CD | D083_TO3_AD
287.026 | 301.9346 | 285.6855 | 275.5296 | 281.5865
5.4395 | 20.3482 | 4.0991 | -6.0568
With Langley Plot Calibration
Langley Plot Adjustments and Lamp Test results included
D083_TO3_C* | D083_TO3_D* | D083_TO3_A* | D083_TO3_CD* | D083_TO3_AD*
284.6259 | 282.826 | 283.9943 | 286.0495 | 284.256
0.36991 | -1.43 | -0.26166 | 1.7935

## IdD083
Dobson 083 at IZO 2016.Sep.14 Langley Plot Results
SLC083Na | -0.9 | SLC083Nc | -0.7 | [9 empty] | SLC083Nd | -0.7 | [11 empty] | Correction to
D083 LP adj. A:-0.70 C:-0.40 D:-0.91 Nad:0.22 | [12 empty] | MaxMu | 2.5 | MaxMu can be changed | [5 empty] | LPC_C_83 |  | LPC_D_83 |  | LPC_A_83 | AD | CD
[20 empty] | With out High mu values | -0.35426 |  | -1.2446 |  | -0.53816 | 0.70647 | 0.89036 | [14 empty] | Limited Mu range
[11 empty] | 48 |  |  |  | 287.026 | 301.9346 | 285.6855 | 275.5296 | 281.5865 | all Values | -0.40351 |  | -0.9136 |  | -0.69599 | 0.21761 | 0.51009 | [14 empty] | 0.88755 |  | 0.70955
D083RC | D083RD | D083RA | day | h | m | s | m | year |  |  |  | D083NC | D083ND | D083NA | D083_TO3_C | D083_TO3_D | D083_TO3_A | D083_TO3_CD | D083_TO3_AD | 1/mu.C | P*C_83 | 1/Mu.D | P*D_83 | 1/Mu.A | P*A_83 | lst.C | Mu.C | AirM.C | lst.D | Mu.D | AirM.D | lst.A | Mu.A | AirM.A | lst.CD | Mu.CD | AirM.CD | lst.AD | Mu.AD | AirM.AD | 1/mu.cd | P*CD_83 | 1/mu.ad | P*AD_83
137.9 | 78.4 | 246.8 | 14 | 7 | 56 | 0 | 9 | 16 | C | 83 | IZO | 129.1168 | 74.8607 | 233.8417 | 285.5909 | 296.8644 | 283.7883 | 277.1027 | 280.4215 | 0.24832 | 0.23421 | 0.2518 | 0.10616 | 0.25529 | 0.50683 | 2016-09-14 07:56 | 4.0271 | 4.201 | 2016-09-14 07:57 | 3.9714 | 4.1374 | 2016-09-14 07:58 | 3.9171 | 4.0757 | 2016-09-14 07:56 | 3.9991 | 4.1689 | 2016-09-14 07:57 | 3.9441 | 4.1063 | 0.25006 | 0.12805 | 0.25354 | 0.40067
130.4 | 74.1 | 235.2 | 14 | 8 | 0 | 0 | 9 | 16 | C | 83 | IZO | 122.0587 | 70.9223 | 221.9193 | 285.5079 | 297.9123 | 284.3328 | 276.1062 | 280.8522 | 0.26227 | 0.23414 | 0.26576 | 0.10654 | 0.26926 | 0.50781 | 2016-09-14 08:00 | 3.8129 | 3.9578 | 2016-09-14 08:01 | 3.7627 | 3.9015 | 2016-09-14 08:02 | 3.7139 | 3.8467 | 2016-09-14 08:01 | 3.7876 | 3.9294 | 2016-09-14 08:02 | 3.7382 | 3.8739 | 0.26402 | 0.1276 | 0.26751 | 0.40127
123.8 | 69.8 | 224.6 | 14 | 8 | 4 | 0 | 9 | 16 | C | 83 | IZO | 115.9024 | 66.9954 | 211.1086 | 286.0143 | 295.8418 | 284.7592 | 278.5739 | 281.9134 | 0.27625 | 0.23456 | 0.27975 | 0.1058 | 0.28324 | 0.50857 | 2016-09-14 08:04 | 3.6199 | 3.7419 | 2016-09-14 08:05 | 3.5747 | 3.6916 | 2016-09-14 08:06 | 3.5305 | 3.6428 | 2016-09-14 08:04 | 3.5971 | 3.7166 | 2016-09-14 08:06 | 3.5525 | 3.667 | 0.278 | 0.12876 | 0.2815 | 0.40277
117.5 | 66.3 | 214.4 | 14 | 8 | 8 | 0 | 9 | 16 | C | 83 | IZO | 110.0458 | 63.8042 | 200.9109 | 285.426 | 296.6124 | 284.4832 | 276.9097 | 281.3833 | 0.29024 | 0.23407 | 0.29373 | 0.10607 | 0.29723 | 0.50807 | 2016-09-14 08:08 | 3.4455 | 3.5489 | 2016-09-14 08:09 | 3.4044 | 3.5038 | 2016-09-14 08:10 | 3.3644 | 3.46 | 2016-09-14 08:08 | 3.4248 | 3.5262 | 2016-09-14 08:10 | 3.3843 | 3.4818 | 0.29199 | 0.128 | 0.29548 | 0.402
111.8 | 63 | 205.1 | 14 | 8 | 12 | 0 | 9 | 16 | C | 83 | IZO | 104.767 | 60.7788 | 191.6826 | 284.9192 | 296.1814 | 284.2597 | 276.3238 | 281.2182 | 0.30422 | 0.23366 | 0.30771 | 0.10592 | 0.3112 | 0.50767 | 2016-09-14 08:12 | 3.2871 | 3.3755 | 2016-09-14 08:13 | 3.2498 | 3.3349 | 2016-09-14 08:14 | 3.2134 | 3.2953 | 2016-09-14 08:12 | 3.2684 | 3.3551 | 2016-09-14 08:13 | 3.2315 | 3.315 | 0.30596 | 0.12774 | 0.30945 | 0.40175
106.9 | 60.1 | 196.8 | 14 | 8 | 16 | 0 | 9 | 16 | C | 83 | IZO | 100.2264 | 58.1083 | 183.5137 | 285.4066 | 296.4456 | 284.4731 | 276.9671 | 281.4241 | 0.31817 | 0.23406 | 0.32166 | 0.10601 | 0.32514 | 0.50806 | 2016-09-14 08:16 | 3.1429 | 3.219 | 2016-09-14 08:17 | 3.1089 | 3.1822 | 2016-09-14 08:18 | 3.0756 | 3.1463 | 2016-09-14 08:17 | 3.1258 | 3.2005 | 2016-09-14 08:18 | 3.0921 | 3.1641 | 0.31992 | 0.12804 | 0.3234 | 0.40204
102.4 | 57.4 | 189.1 | 14 | 8 | 20 | 0 | 9 | 16 | C | 83 | IZO | 96.0484 | 55.6082 | 175.9253 | 285.7404 | 296.0931 | 284.4586 | 277.8176 | 281.4987 | 0.3321 | 0.23433 | 0.33557 | 0.10589 | 0.33904 | 0.50803 | 2016-09-14 08:20 | 3.0112 | 3.077 | 2016-09-14 08:21 | 2.98 | 3.0435 | 2016-09-14 08:22 | 2.9495 | 3.0108 | 2016-09-14 08:20 | 2.9955 | 3.0601 | 2016-09-14 08:22 | 2.9646 | 3.0271 | 0.33384 | 0.12844 | 0.33731 | 0.40214
98.3 | 55.1 | 182.5 | 14 | 8 | 24 | 0 | 9 | 16 | C | 83 | IZO | 92.2357 | 53.4693 | 169.4169 | 286.1165 | 297.1911 | 285.3366 | 277.6239 | 282.3252 | 0.34598 | 0.23464 | 0.34944 | 0.10628 | 0.3529 | 0.5096 | 2016-09-14 08:24 | 2.8904 | 2.9476 | 2016-09-14 08:25 | 2.8617 | 2.9171 | 2016-09-14 08:26 | 2.8337 | 2.8872 | 2016-09-14 08:24 | 2.876 | 2.9323 | 2016-09-14 08:26 | 2.8476 | 2.902 | 0.34771 | 0.12836 | 0.35117 | 0.40332
94.4 | 53.1 | 176 | 14 | 8 | 28 | 0 | 9 | 16 | C | 83 | IZO | 88.6195 | 51.5978 | 163.0364 | 285.9803 | 299.212 | 285.4022 | 275.81 | 281.9045 | 0.3598 | 0.23453 | 0.36325 | 0.107 | 0.36669 | 0.50972 | 2016-09-14 08:28 | 2.7793 | 2.8294 | 2016-09-14 08:29 | 2.7529 | 2.8014 | 2016-09-14 08:30 | 2.7271 | 2.774 | 2016-09-14 08:28 | 2.7661 | 2.8153 | 2016-09-14 08:29 | 2.74 | 2.7876 | 0.36152 | 0.12753 | 0.36497 | 0.40271
90.9 | 51 | 169.7 | 14 | 8 | 32 | 0 | 9 | 16 | C | 83 | IZO | 85.3767 | 49.618 | 156.8816 | 286.2384 | 298.5107 | 284.8942 | 276.8014 | 281.4475 | 0.37356 | 0.23474 | 0.37699 | 0.10675 | 0.38041 | 0.50881 | 2016-09-14 08:32 | 2.6769 | 2.7209 | 2016-09-14 08:33 | 2.6526 | 2.6952 | 2016-09-14 08:34 | 2.6287 | 2.67 | 2016-09-14 08:33 | 2.6647 | 2.708 | 2016-09-14 08:34 | 2.6406 | 2.6825 | 0.37528 | 0.12799 | 0.3787 | 0.40205
87.7 | 49.1 | 164.2 | 14 | 8 | 36 | 0 | 9 | 16 | C | 83 | IZO | 82.4083 | 47.8156 | 151.4999 | 286.6137 | 298.0886 | 285.0742 | 277.7863 | 281.7804 | 0.38725 | 0.23505 | 0.39066 | 0.1066 | 0.39406 | 0.50913 | 2016-09-14 08:36 | 2.5823 | 2.6211 | 2016-09-14 08:37 | 2.5598 | 2.5974 | 2016-09-14 08:38 | 2.5377 | 2.5741 | 2016-09-14 08:36 | 2.571 | 2.6092 | 2016-09-14 08:38 | 2.5487 | 2.5857 | 0.38895 | 0.12845 | 0.39236 | 0.40253
84.5 | 47.5 | 159.2 | 14 | 8 | 40 | 0 | 9 | 16 | C | 83 | IZO | 79.4477 | 46.289 | 146.624 | 285.9444 | 299.2688 | 285.511 | 275.6808 | 282.0331 | 0.40086 | 0.2345 | 0.40425 | 0.10702 | 0.40763 | 0.50991 | 2016-09-14 08:40 | 2.4946 | 2.529 | 2016-09-14 08:41 | 2.4737 | 2.5071 | 2016-09-14 08:42 | 2.4532 | 2.4856 | 2016-09-14 08:40 | 2.4841 | 2.518 | 2016-09-14 08:42 | 2.4634 | 2.4963 | 0.40255 | 0.12748 | 0.40594 | 0.40289
81.9 | 45.9 | 154.4 | 14 | 8 | 44 | 0 | 9 | 16 | C | 83 | IZO | 77.0538 | 44.7538 | 141.979 | 287.0613 | 299.05 | 285.674 | 277.8251 | 282.2934 | 0.41438 | 0.23542 | 0.41775 | 0.10695 | 0.42111 | 0.5102 | 2016-09-14 08:44 | 2.4132 | 2.4438 | 2016-09-14 08:45 | 2.3938 | 2.4234 | 2016-09-14 08:46 | 2.3747 | 2.4035 | 2016-09-14 08:44 | 2.4035 | 2.4336 | 2016-09-14 08:46 | 2.3842 | 2.4134 | 0.41606 | 0.12847 | 0.41943 | 0.40325
79.2 | 44.5 | 149.8 | 14 | 8 | 48 | 0 | 9 | 16 | C | 83 | IZO | 74.5708 | 43.4039 | 137.5175 | 286.8277 | 299.7665 | 285.5032 | 276.8524 | 281.9018 | 0.42781 | 0.23522 | 0.43115 | 0.1072 | 0.43449 | 0.5099 | 2016-09-14 08:48 | 2.3375 | 2.3647 | 2016-09-14 08:49 | 2.3194 | 2.3458 | 2016-09-14 08:50 | 2.3016 | 2.3273 | 2016-09-14 08:49 | 2.3284 | 2.3552 | 2016-09-14 08:50 | 2.3104 | 2.3365 | 0.42948 | 0.12802 | 0.43282 | 0.40269
76.8 | 42.9 | 145.7 | 14 | 8 | 52 | 0 | 9 | 16 | C | 83 | IZO | 72.3603 | 41.8534 | 133.5776 | 287.1489 | 296.838 | 285.8907 | 279.6824 | 283.1208 | 0.44114 | 0.23549 | 0.44446 | 0.10615 | 0.44776 | 0.51059 | 2016-09-14 08:52 | 2.2669 | 2.2912 | 2016-09-14 08:53 | 2.2499 | 2.2736 | 2016-09-14 08:54 | 2.2333 | 2.2564 | 2016-09-14 08:52 | 2.2584 | 2.2824 | 2016-09-14 08:54 | 2.2416 | 2.265 | 0.4428 | 0.12933 | 0.44611 | 0.40443
74.4 | 41.8 | 141.7 | 14 | 8 | 56 | 0 | 9 | 16 | C | 83 | IZO | 70.1486 | 40.7823 | 129.7576 | 286.6458 | 298.7524 | 285.9187 | 277.3066 | 282.6777 | 0.45436 | 0.23507 | 0.45765 | 0.10684 | 0.46093 | 0.51064 | 2016-09-14 08:56 | 2.2009 | 2.2227 | 2016-09-14 08:57 | 2.1851 | 2.2063 | 2016-09-14 08:58 | 2.1695 | 2.1902 | 2016-09-14 08:56 | 2.1929 | 2.2145 | 2016-09-14 08:58 | 2.1773 | 2.1982 | 0.45601 | 0.12824 | 0.45929 | 0.4038
72.3 | 40.5 | 137.9 | 14 | 9 | 0 | 0 | 9 | 16 | C | 83 | IZO | 68.2133 | 39.5109 | 126.1359 | 286.9052 | 297.097 | 285.8258 | 279.0434 | 282.977 | 0.46748 | 0.23529 | 0.47074 | 0.10625 | 0.47399 | 0.51047 | 2016-09-14 09:00 | 2.1392 | 2.1588 | 2016-09-14 09:01 | 2.1243 | 2.1435 | 2016-09-14 09:02 | 2.1098 | 2.1284 | 2016-09-14 09:00 | 2.1317 | 2.1511 | 2016-09-14 09:02 | 2.117 | 2.1359 | 0.46911 | 0.12904 | 0.47236 | 0.40422
70.5 | 39.3 | 134.4 | 14 | 9 | 4 | 0 | 9 | 16 | C | 83 | IZO | 66.5555 | 38.3315 | 122.7964 | 288.0699 | 295.6148 | 285.8875 | 282.252 | 283.4261 | 0.48047 | 0.23624 | 0.4837 | 0.10572 | 0.48692 | 0.51058 | 2016-09-14 09:04 | 2.0813 | 2.099 | 2016-09-14 09:05 | 2.0674 | 2.0846 | 2016-09-14 09:06 | 2.0537 | 2.0705 | 2016-09-14 09:05 | 2.0743 | 2.0918 | 2016-09-14 09:06 | 2.0605 | 2.0776 | 0.48209 | 0.13053 | 0.48531 | 0.40486
68.4 | 38.5 | 131 | 14 | 9 | 8 | 0 | 9 | 16 | C | 83 | IZO | 64.6224 | 37.5419 | 119.5461 | 286.9506 | 298.5496 | 285.629 | 277.9956 | 282.3689 | 0.49334 | 0.23532 | 0.49654 | 0.10677 | 0.49973 | 0.51012 | 2016-09-14 09:08 | 2.027 | 2.043 | 2016-09-14 09:09 | 2.0139 | 2.0295 | 2016-09-14 09:10 | 2.0011 | 2.0163 | 2016-09-14 09:08 | 2.0204 | 2.0362 | 2016-09-14 09:10 | 2.0075 | 2.0229 | 0.49494 | 0.12856 | 0.49814 | 0.40335
66.7 | 37.4 | 128 | 14 | 9 | 12 | 0 | 9 | 16 | C | 83 | IZO | 63.0562 | 36.4523 | 116.6843 | 287.3853 | 296.5012 | 285.9316 | 280.3482 | 283.2608 | 0.50609 | 0.23568 | 0.50925 | 0.10603 | 0.51241 | 0.51066 | 2016-09-14 09:12 | 1.9759 | 1.9904 | 2016-09-14 09:13 | 1.9637 | 1.9778 | 2016-09-14 09:14 | 1.9516 | 1.9653 | 2016-09-14 09:12 | 1.9698 | 1.9841 | 2016-09-14 09:14 | 1.9576 | 1.9715 | 0.50767 | 0.12965 | 0.51083 | 0.40463
55.4 | 31.7 | 107.9 | 14 | 9 | 45 | 0 | 9 | 16 | C | 83 | IZO | 52.5677 | 30.7508 | 97.7697 | 286.8747 | 301.4467 | 286.0333 | 275.6075 | 282.1533 | 0.60576 | 0.23526 | 0.60861 | 0.1078 | 0.61145 | 0.51084 | 2016-09-14 09:45 | 1.6508 | 1.6576 | 2016-09-14 09:46 | 1.6431 | 1.6497 | 2016-09-14 09:47 | 1.6355 | 1.6419 | 2016-09-14 09:46 | 1.6469 | 1.6536 | 2016-09-14 09:47 | 1.6393 | 1.6458 | 0.60719 | 0.12746 | 0.61003 | 0.40304
54.5 | 30.8 | 105.3 | 14 | 9 | 50 | 0 | 9 | 16 | C | 83 | IZO | 51.7251 | 29.8431 | 95.3514 | 289.6143 | 297.8424 | 285.2711 | 283.2514 | 282.1031 | 0.61992 | 0.23751 | 0.62271 | 0.10651 | 0.6255 | 0.50948 | 2016-09-14 09:50 | 1.6131 | 1.6192 | 2016-09-14 09:51 | 1.6059 | 1.6118 | 2016-09-14 09:52 | 1.5987 | 1.6045 | 2016-09-14 09:50 | 1.6095 | 1.6155 | 2016-09-14 09:52 | 1.6023 | 1.6081 | 0.62132 | 0.131 | 0.62411 | 0.40297
52.9 | 30.2 | 102.9 | 14 | 9 | 55 | 0 | 9 | 16 | C | 83 | IZO | 50.2217 | 29.2366 | 93.1078 | 286.7852 | 298.6569 | 284.6098 | 277.604 | 281.0726 | 0.6338 | 0.23519 | 0.63655 | 0.10681 | 0.63928 | 0.5083 | 2016-09-14 09:55 | 1.5778 | 1.5832 | 2016-09-14 09:56 | 1.571 | 1.5763 | 2016-09-14 09:57 | 1.5643 | 1.5694 | 2016-09-14 09:56 | 1.5744 | 1.5797 | 2016-09-14 09:56 | 1.5676 | 1.5729 | 0.63518 | 0.12838 | 0.63791 | 0.40149
51.7 | 30 | 101.6 | 14 | 9 | 59 | 0 | 9 | 16 | C | 83 | IZO | 49.0894 | 29.0342 | 91.8876 | 284.588 | 303.9241 | 285.8191 | 269.6344 | 281.2662 | 0.64471 | 0.23339 | 0.64741 | 0.10869 | 0.6501 | 0.51046 | 2016-09-14 09:59 | 1.5511 | 1.556 | 2016-09-14 10:00 | 1.5446 | 1.5495 | 2016-09-14 10:01 | 1.5382 | 1.543 | 2016-09-14 09:59 | 1.5478 | 1.5528 | 2016-09-14 10:00 | 1.5414 | 1.5462 | 0.64606 | 0.1247 | 0.64876 | 0.40177
51.1 | 29.4 | 99.9 | 14 | 10 | 3 | 0 | 9 | 16 | C | 83 | IZO | 48.5224 | 28.4262 | 90.2892 | 286.499 | 301.4533 | 285.3876 | 274.933 | 281.3454 | 0.65544 | 0.23495 | 0.65809 | 0.10781 | 0.66073 | 0.50969 | 2016-09-14 10:03 | 1.5257 | 1.5303 | 2016-09-14 10:04 | 1.5195 | 1.524 | 2016-09-14 10:05 | 1.5135 | 1.5178 | 2016-09-14 10:04 | 1.5226 | 1.5271 | 2016-09-14 10:04 | 1.5165 | 1.5209 | 0.65677 | 0.12715 | 0.65941 | 0.40188
50.2 | 29.1 | 98.6 | 14 | 10 | 7 | 0 | 9 | 16 | C | 83 | IZO | 47.6703 | 28.1217 | 89.0662 | 285.8376 | 304.1465 | 286.0827 | 271.677 | 281.5404 | 0.66598 | 0.23441 | 0.66859 | 0.10877 | 0.67118 | 0.51093 | 2016-09-14 10:07 | 1.5015 | 1.5057 | 2016-09-14 10:08 | 1.4957 | 1.4998 | 2016-09-14 10:09 | 1.4899 | 1.4939 | 2016-09-14 10:08 | 1.4986 | 1.5028 | 2016-09-14 10:08 | 1.4928 | 1.4968 | 0.66729 | 0.12564 | 0.66989 | 0.40216
49.5 | 28.7 | 97 | 14 | 10 | 11 | 0 | 9 | 16 | C | 83 | IZO | 47.0062 | 27.7152 | 87.5615 | 286.3976 | 304.5743 | 285.4658 | 272.3387 | 280.6621 | 0.67633 | 0.23487 | 0.67889 | 0.10892 | 0.68144 | 0.50983 | 2016-09-14 10:11 | 1.4786 | 1.4824 | 2016-09-14 10:12 | 1.473 | 1.4768 | 2016-09-14 10:13 | 1.4675 | 1.4712 | 2016-09-14 10:12 | 1.4758 | 1.4796 | 2016-09-14 10:12 | 1.4702 | 1.474 | 0.67762 | 0.12595 | 0.68017 | 0.40091
48.7 | 28.3 | 95.8 | 14 | 10 | 15 | 0 | 9 | 16 | C | 83 | IZO | 46.2454 | 27.3081 | 86.4316 | 285.8704 | 304.6065 | 286.0333 | 271.3786 | 281.3638 | 0.6865 | 0.23444 | 0.68901 | 0.10893 | 0.6915 | 0.51084 | 2016-09-14 10:15 | 1.4567 | 1.4602 | 2016-09-14 10:16 | 1.4514 | 1.4548 | 2016-09-14 10:17 | 1.4461 | 1.4495 | 2016-09-14 10:16 | 1.454 | 1.4575 | 2016-09-14 10:16 | 1.4487 | 1.4522 | 0.68775 | 0.12551 | 0.69026 | 0.40191
48.2 | 28 | 94.4 | 14 | 10 | 19 | 0 | 9 | 16 | C | 83 | IZO | 45.7687 | 27.0024 | 85.1124 | 287.4586 | 306.2393 | 285.6351 | 272.9317 | 280.4573 | 0.69646 | 0.23574 | 0.69892 | 0.10952 | 0.70137 | 0.51013 | 2016-09-14 10:19 | 1.4358 | 1.4391 | 2016-09-14 10:20 | 1.4308 | 1.434 | 2016-09-14 10:21 | 1.4258 | 1.4289 | 2016-09-14 10:20 | 1.4333 | 1.4365 | 2016-09-14 10:20 | 1.4283 | 1.4314 | 0.69769 | 0.12622 | 0.70015 | 0.40061
47.4 | 27.6 | 93.2 | 14 | 10 | 23 | 0 | 9 | 16 | C | 83 | IZO | 45.0041 | 26.5942 | 83.9826 | 286.3398 | 305.5213 | 285.7506 | 271.5026 | 280.7816 | 0.70623 | 0.23482 | 0.70864 | 0.10926 | 0.71103 | 0.51034 | 2016-09-14 10:23 | 1.416 | 1.419 | 2016-09-14 10:24 | 1.4112 | 1.4141 | 2016-09-14 10:25 | 1.4064 | 1.4093 | 2016-09-14 10:24 | 1.4136 | 1.4165 | 2016-09-14 10:24 | 1.4088 | 1.4117 | 0.70743 | 0.12556 | 0.70984 | 0.40108
47.3 | 27.2 | 92.2 | 14 | 10 | 27 | 0 | 9 | 16 | C | 83 | IZO | 44.9083 | 26.1855 | 83.043 | 290.7671 | 304.4191 | 286.4161 | 280.205 | 281.8898 | 0.71579 | 0.23845 | 0.71815 | 0.10887 | 0.7205 | 0.51153 | 2016-09-14 10:27 | 1.3971 | 1.3998 | 2016-09-14 10:28 | 1.3925 | 1.3952 | 2016-09-14 10:29 | 1.3879 | 1.3906 | 2016-09-14 10:28 | 1.3948 | 1.3975 | 2016-09-14 10:28 | 1.3902 | 1.3928 | 0.71697 | 0.12959 | 0.71933 | 0.40266
46.3 | 26.8 | 90.9 | 14 | 10 | 31 | 0 | 9 | 16 | C | 83 | IZO | 43.9481 | 25.7762 | 81.8257 | 287.4156 | 302.9375 | 285.7519 | 275.4074 | 281.4303 | 0.72515 | 0.23571 | 0.72746 | 0.10834 | 0.72975 | 0.51034 | 2016-09-14 10:31 | 1.379 | 1.3815 | 2016-09-14 10:32 | 1.3747 | 1.3771 | 2016-09-14 10:33 | 1.3703 | 1.3727 | 2016-09-14 10:32 | 1.3768 | 1.3793 | 2016-09-14 10:32 | 1.3725 | 1.3749 | 0.72631 | 0.12737 | 0.72861 | 0.402
45.6 | 26.6 | 89.8 | 14 | 10 | 35 | 0 | 9 | 16 | C | 83 | IZO | 43.2735 | 25.5713 | 80.7993 | 286.2918 | 305.3036 | 285.6561 | 271.5847 | 280.7183 | 0.7343 | 0.23478 | 0.73656 | 0.10918 | 0.7388 | 0.51017 | 2016-09-14 10:35 | 1.3618 | 1.3642 | 2016-09-14 10:36 | 1.3577 | 1.3599 | 2016-09-14 10:37 | 1.3536 | 1.3558 | 2016-09-14 10:36 | 1.3597 | 1.362 | 2016-09-14 10:36 | 1.3556 | 1.3578 | 0.73543 | 0.1256 | 0.73768 | 0.40099
45.3 | 26.4 | 88.8 | 14 | 10 | 39 | 0 | 9 | 16 | C | 83 | IZO | 42.9839 | 25.3663 | 79.8683 | 288.3962 | 307.4105 | 285.7607 | 273.6868 | 280.3217 | 0.74324 | 0.23651 | 0.74544 | 0.10994 | 0.74763 | 0.51036 | 2016-09-14 10:39 | 1.3455 | 1.3476 | 2016-09-14 10:40 | 1.3415 | 1.3436 | 2016-09-14 10:41 | 1.3376 | 1.3396 | 2016-09-14 10:40 | 1.3435 | 1.3456 | 2016-09-14 10:40 | 1.3395 | 1.3416 | 0.74434 | 0.12657 | 0.74654 | 0.40042
44.7 | 26.2 | 87.9 | 14 | 10 | 43 | 0 | 9 | 16 | C | 83 | IZO | 42.4038 | 25.1612 | 79.032 | 287.6613 | 309.2599 | 286.078 | 270.9534 | 280.2555 | 0.75197 | 0.23591 | 0.75411 | 0.1106 | 0.75625 | 0.51092 | 2016-09-14 10:43 | 1.3298 | 1.3318 | 2016-09-14 10:44 | 1.3261 | 1.328 | 2016-09-14 10:45 | 1.3223 | 1.3242 | 2016-09-14 10:44 | 1.3279 | 1.3299 | 2016-09-14 10:44 | 1.3242 | 1.3261 | 0.75304 | 0.12531 | 0.75518 | 0.40032
44.2 | 25.8 | 86.9 | 14 | 10 | 47 | 0 | 9 | 16 | C | 83 | IZO | 41.9193 | 24.7505 | 78.1039 | 287.5792 | 306.4735 | 285.8262 | 272.9621 | 280.6384 | 0.76047 | 0.23584 | 0.76257 | 0.1096 | 0.76464 | 0.51047 | 2016-09-14 10:47 | 1.315 | 1.3168 | 2016-09-14 10:48 | 1.3114 | 1.3131 | 2016-09-14 10:49 | 1.3078 | 1.3095 | 2016-09-14 10:48 | 1.3132 | 1.3149 | 2016-09-14 10:48 | 1.3096 | 1.3113 | 0.76152 | 0.12624 | 0.76361 | 0.40087
43.9 | 25.5 | 86.2 | 14 | 10 | 51 | 0 | 9 | 16 | C | 83 | IZO | 41.6283 | 24.4421 | 77.4549 | 289.0972 | 305.5469 | 286.5987 | 276.37 | 281.8363 | 0.76876 | 0.23709 | 0.7708 | 0.10927 | 0.77282 | 0.51185 | 2016-09-14 10:51 | 1.3008 | 1.3024 | 2016-09-14 10:52 | 1.2974 | 1.299 | 2016-09-14 10:53 | 1.294 | 1.2955 | 2016-09-14 10:52 | 1.2991 | 1.3007 | 2016-09-14 10:52 | 1.2957 | 1.2972 | 0.76978 | 0.12782 | 0.77181 | 0.40258
43.4 | 25.3 | 85.4 | 14 | 10 | 55 | 0 | 9 | 16 | C | 83 | IZO | 41.1423 | 24.2364 | 76.7136 | 288.5982 | 306.5615 | 286.8117 | 274.7004 | 281.8487 | 0.77683 | 0.23668 | 0.77881 | 0.10963 | 0.78078 | 0.51223 | 2016-09-14 10:55 | 1.2873 | 1.2888 | 2016-09-14 10:56 | 1.284 | 1.2855 | 2016-09-14 10:57 | 1.2808 | 1.2822 | 2016-09-14 10:56 | 1.2856 | 1.2871 | 2016-09-14 10:56 | 1.2824 | 1.2838 | 0.77782 | 0.12704 | 0.77979 | 0.4026
42.9 | 25.2 | 84.8 | 14 | 10 | 59 | 0 | 9 | 16 | C | 83 | IZO | 40.6554 | 24.1335 | 76.1578 | 287.8821 | 309.5927 | 287.6813 | 271.0864 | 282.1771 | 0.78467 | 0.23609 | 0.78659 | 0.11072 | 0.7885 | 0.51379 | 2016-09-14 10:59 | 1.2744 | 1.2758 | 2016-09-14 11:00 | 1.2713 | 1.2726 | 2016-09-14 11:01 | 1.2682 | 1.2695 | 2016-09-14 11:00 | 1.2729 | 1.2742 | 2016-09-14 11:01 | 1.2698 | 1.2711 | 0.78563 | 0.12537 | 0.78755 | 0.40307
42.9 | 25.1 | 84.1 | 14 | 11 | 3 | 0 | 9 | 16 | C | 83 | IZO | 40.6554 | 24.0305 | 75.51 | 291.6649 | 312.424 | 287.9957 | 275.6046 | 281.8615 | 0.79228 | 0.23919 | 0.79415 | 0.11173 | 0.796 | 0.51435 | 2016-09-14 11:03 | 1.2622 | 1.2634 | 2016-09-14 11:04 | 1.2592 | 1.2604 | 2016-09-14 11:05 | 1.2563 | 1.2575 | 2016-09-14 11:04 | 1.2607 | 1.2619 | 2016-09-14 11:04 | 1.2577 | 1.2589 | 0.79322 | 0.12746 | 0.79508 | 0.40262
42.2 | 25.1 | 83.4 | 14 | 11 | 7 | 0 | 9 | 16 | C | 83 | IZO | 39.972 | 24.0305 | 74.8633 | 288.6693 | 317.3639 | 288.1636 | 266.4729 | 280.8347 | 0.79967 | 0.23673 | 0.80147 | 0.1135 | 0.80327 | 0.51465 | 2016-09-14 11:07 | 1.2505 | 1.2517 | 2016-09-14 11:08 | 1.2477 | 1.2488 | 2016-09-14 11:09 | 1.2449 | 1.246 | 2016-09-14 11:08 | 1.2491 | 1.2502 | 2016-09-14 11:08 | 1.2463 | 1.2474 | 0.80057 | 0.12324 | 0.80237 | 0.40115
41.8 | 24.8 | 82.4 | 14 | 11 | 11 | 0 | 9 | 16 | C | 83 | IZO | 39.5805 | 23.7215 | 73.9408 | 288.3112 | 315.1587 | 286.9302 | 267.5428 | 279.8447 | 0.80682 | 0.23644 | 0.80857 | 0.11271 | 0.8103 | 0.51244 | 2016-09-14 11:11 | 1.2394 | 1.2405 | 2016-09-14 11:12 | 1.2368 | 1.2378 | 2016-09-14 11:13 | 1.2341 | 1.2351 | 2016-09-14 11:12 | 1.2381 | 1.2391 | 2016-09-14 11:12 | 1.2354 | 1.2364 | 0.80769 | 0.12373 | 0.80944 | 0.39974
Note on B46: Bob Evans: Edited from 25.7

## IdD083_star
This Page uses the lamp tests and the langley plot results in the limited mu range from sheet IdD083 | [21 empty] | the correction values from ID083 are applied to the NC
Dobson 083 at IZO 2016.Sep.14 Langley Plot Results | [15 empty] | 109
SLC083Na | -0.9 | SLC083Nc | -0.7 | [9 empty] | SLC083Nd | -0.7
D083 LP adj. A:-0.16 C:-0.05 D:0.25 Nad:-0.41 | [20 empty] | LPC_C_83 |  | LPC_D_83 |  | LPC_A_83
[15 empty] | 284.6259 | 282.826 | 283.9943 | 286.0495 | 284.256 | All Obs | -0.049248 |  | 0.24739 |  | -0.15784 | -0.40523 | -0.29664
D083RC | D083RD | D083RA | day | h | m | s | m | year |  |  |  | D083NC | D083ND | D083NA | D083_TO3_C* | D083_TO3_D* | D083_TO3_A* | D083_TO3_CD* | D083_TO3_AD* | 1/mu.C | P*C_83 | 1/Mu.D | P*D_83 | 1/Mu.A | P*A_83 | lst.C | Mu.C | AirM.C | lst.D | Mu.D | AirM.D | lst.A | Mu.A | AirM.A | lst.CD | Mu.CD | AirM.CD | lst.AD | Mu.AD | AirM.AD
137.9 | 78.4 | 246.8 | 14 | 7 | 56 | 0 | 9 | 16 | C | 83 | IZO | 128.7625 | 73.616 | 233.3036 | 284.5182 | 288.101 | 283.0191 | 281.977 | 281.6537 | 0.24832 | 0.23333 | 0.2518 | 0.10303 | 0.25529 | 0.50546 | 2016-09-14 07:56 | 4.0271 | 4.201 | 2016-09-14 07:57 | 3.9714 | 4.1374 | 2016-09-14 07:58 | 3.9171 | 4.0757 | 2016-09-14 07:56 | 3.9991 | 4.1689 | 2016-09-14 07:57 | 3.9441 | 4.1063
130.4 | 74.1 | 235.2 | 14 | 8 | 0 | 0 | 9 | 16 | C | 83 | IZO | 121.7044 | 69.6777 | 221.3811 | 284.375 | 288.6629 | 283.5215 | 281.2493 | 282.1534 | 0.26227 | 0.23321 | 0.26576 | 0.10323 | 0.26926 | 0.50636 | 2016-09-14 08:00 | 3.8129 | 3.9578 | 2016-09-14 08:01 | 3.7627 | 3.9015 | 2016-09-14 08:02 | 3.7139 | 3.8467 | 2016-09-14 08:01 | 3.7876 | 3.9294 | 2016-09-14 08:02 | 3.7382 | 3.8739
123.8 | 69.8 | 224.6 | 14 | 8 | 4 | 0 | 9 | 16 | C | 83 | IZO | 115.5481 | 65.7508 | 210.5705 | 284.821 | 286.1058 | 283.9057 | 283.9863 | 283.2838 | 0.27625 | 0.23358 | 0.27975 | 0.10232 | 0.28324 | 0.50704 | 2016-09-14 08:04 | 3.6199 | 3.7419 | 2016-09-14 08:05 | 3.5747 | 3.6916 | 2016-09-14 08:06 | 3.5305 | 3.6428 | 2016-09-14 08:04 | 3.5971 | 3.7166 | 2016-09-14 08:06 | 3.5525 | 3.667
117.5 | 66.3 | 214.4 | 14 | 8 | 8 | 0 | 9 | 16 | C | 83 | IZO | 109.6915 | 62.5596 | 200.3728 | 284.1722 | 286.3895 | 283.5876 | 282.5914 | 282.8229 | 0.29024 | 0.23305 | 0.29373 | 0.10242 | 0.29723 | 0.50647 | 2016-09-14 08:08 | 3.4455 | 3.5489 | 2016-09-14 08:09 | 3.4044 | 3.5038 | 2016-09-14 08:10 | 3.3644 | 3.46 | 2016-09-14 08:08 | 3.4248 | 3.5262 | 2016-09-14 08:10 | 3.3843 | 3.4818
111.8 | 63 | 205.1 | 14 | 8 | 12 | 0 | 9 | 16 | C | 83 | IZO | 104.4127 | 59.5342 | 191.1444 | 283.605 | 285.4722 | 283.322 | 282.2745 | 282.7269 | 0.30422 | 0.23258 | 0.30771 | 0.10209 | 0.3112 | 0.506 | 2016-09-14 08:12 | 3.2871 | 3.3755 | 2016-09-14 08:13 | 3.2498 | 3.3349 | 2016-09-14 08:14 | 3.2134 | 3.2953 | 2016-09-14 08:12 | 3.2684 | 3.3551 | 2016-09-14 08:13 | 3.2315 | 3.315
106.9 | 60.1 | 196.8 | 14 | 8 | 16 | 0 | 9 | 16 | C | 83 | IZO | 99.8722 | 56.8637 | 182.9755 | 284.0321 | 285.2508 | 283.4934 | 283.1862 | 283.0019 | 0.31817 | 0.23293 | 0.32166 | 0.10201 | 0.32514 | 0.50631 | 2016-09-14 08:16 | 3.1429 | 3.219 | 2016-09-14 08:17 | 3.1089 | 3.1822 | 2016-09-14 08:18 | 3.0756 | 3.1463 | 2016-09-14 08:17 | 3.1258 | 3.2005 | 2016-09-14 08:18 | 3.0921 | 3.1641
102.4 | 57.4 | 189.1 | 14 | 8 | 20 | 0 | 9 | 16 | C | 83 | IZO | 95.6941 | 54.3636 | 175.3871 | 284.3058 | 284.4142 | 283.437 | 284.3045 | 283.1453 | 0.3321 | 0.23316 | 0.33557 | 0.10171 | 0.33904 | 0.50621 | 2016-09-14 08:20 | 3.0112 | 3.077 | 2016-09-14 08:21 | 2.98 | 3.0435 | 2016-09-14 08:22 | 2.9495 | 3.0108 | 2016-09-14 08:20 | 2.9955 | 3.0601 | 2016-09-14 08:22 | 2.9646 | 3.0271
98.3 | 55.1 | 182.5 | 14 | 8 | 24 | 0 | 9 | 16 | C | 83 | IZO | 91.8814 | 52.2246 | 168.8788 | 284.6219 | 285.0296 | 284.2732 | 284.3776 | 284.0404 | 0.34598 | 0.23342 | 0.34944 | 0.10193 | 0.3529 | 0.5077 | 2016-09-14 08:24 | 2.8904 | 2.9476 | 2016-09-14 08:25 | 2.8617 | 2.9171 | 2016-09-14 08:26 | 2.8337 | 2.8872 | 2016-09-14 08:24 | 2.876 | 2.9323 | 2016-09-14 08:26 | 2.8476 | 2.902
94.4 | 53.1 | 176 | 14 | 8 | 28 | 0 | 9 | 16 | C | 83 | IZO | 88.2653 | 50.3531 | 162.4983 | 284.4261 | 286.5699 | 284.2973 | 282.8294 | 283.6881 | 0.3598 | 0.23325 | 0.36325 | 0.10248 | 0.36669 | 0.50774 | 2016-09-14 08:28 | 2.7793 | 2.8294 | 2016-09-14 08:29 | 2.7529 | 2.8014 | 2016-09-14 08:30 | 2.7271 | 2.774 | 2016-09-14 08:28 | 2.7661 | 2.8153 | 2016-09-14 08:29 | 2.74 | 2.7876
90.9 | 51 | 169.7 | 14 | 8 | 32 | 0 | 9 | 16 | C | 83 | IZO | 85.0224 | 48.3734 | 156.3434 | 284.6247 | 285.3903 | 283.7479 | 284.0854 | 283.2991 | 0.37356 | 0.23342 | 0.37699 | 0.10206 | 0.38041 | 0.50676 | 2016-09-14 08:32 | 2.6769 | 2.7209 | 2016-09-14 08:33 | 2.6526 | 2.6952 | 2016-09-14 08:34 | 2.6287 | 2.67 | 2016-09-14 08:33 | 2.6647 | 2.708 | 2016-09-14 08:34 | 2.6406 | 2.6825
87.7 | 49.1 | 164.2 | 14 | 8 | 36 | 0 | 9 | 16 | C | 83 | IZO | 82.054 | 46.571 | 150.9618 | 284.9409 | 284.4925 | 283.8867 | 285.3332 | 283.6997 | 0.38725 | 0.23368 | 0.39066 | 0.10174 | 0.39406 | 0.50701 | 2016-09-14 08:36 | 2.5823 | 2.6211 | 2016-09-14 08:37 | 2.5598 | 2.5974 | 2016-09-14 08:38 | 2.5377 | 2.5741 | 2016-09-14 08:36 | 2.571 | 2.6092 | 2016-09-14 08:38 | 2.5487 | 2.5857
84.5 | 47.5 | 159.2 | 14 | 8 | 40 | 0 | 9 | 16 | C | 83 | IZO | 79.0934 | 45.0444 | 146.0859 | 284.2127 | 285.1997 | 284.2827 | 283.4891 | 284.0197 | 0.40086 | 0.23308 | 0.40425 | 0.10199 | 0.40763 | 0.50772 | 2016-09-14 08:40 | 2.4946 | 2.529 | 2016-09-14 08:41 | 2.4737 | 2.5071 | 2016-09-14 08:42 | 2.4532 | 2.4856 | 2016-09-14 08:40 | 2.4841 | 2.518 | 2016-09-14 08:42 | 2.4634 | 2.4963
81.9 | 45.9 | 154.4 | 14 | 8 | 44 | 0 | 9 | 16 | C | 83 | IZO | 76.6995 | 43.5092 | 141.4408 | 285.2712 | 284.5111 | 284.4051 | 285.8932 | 284.3468 | 0.41438 | 0.23395 | 0.41775 | 0.10175 | 0.42111 | 0.50793 | 2016-09-14 08:44 | 2.4132 | 2.4438 | 2016-09-14 08:45 | 2.3938 | 2.4234 | 2016-09-14 08:46 | 2.3747 | 2.4035 | 2016-09-14 08:44 | 2.4035 | 2.4336 | 2016-09-14 08:46 | 2.3842 | 2.4134
79.2 | 44.5 | 149.8 | 14 | 8 | 48 | 0 | 9 | 16 | C | 83 | IZO | 74.2165 | 42.1593 | 136.9793 | 284.9796 | 284.761 | 284.1939 | 285.1783 | 284.0217 | 0.42781 | 0.23371 | 0.43115 | 0.10184 | 0.43449 | 0.50756 | 2016-09-14 08:48 | 2.3375 | 2.3647 | 2016-09-14 08:49 | 2.3194 | 2.3458 | 2016-09-14 08:50 | 2.3016 | 2.3273 | 2016-09-14 08:49 | 2.3284 | 2.3552 | 2016-09-14 08:50 | 2.3104 | 2.3365
76.8 | 42.9 | 145.7 | 14 | 8 | 52 | 0 | 9 | 16 | C | 83 | IZO | 72.006 | 40.6087 | 133.0395 | 285.2433 | 281.3696 | 284.5414 | 288.2642 | 285.3065 | 0.44114 | 0.23392 | 0.44446 | 0.10062 | 0.44776 | 0.50818 | 2016-09-14 08:52 | 2.2669 | 2.2912 | 2016-09-14 08:53 | 2.2499 | 2.2736 | 2016-09-14 08:54 | 2.2333 | 2.2564 | 2016-09-14 08:52 | 2.2584 | 2.2824 | 2016-09-14 08:54 | 2.2416 | 2.265
74.4 | 41.8 | 141.7 | 14 | 8 | 56 | 0 | 9 | 16 | C | 83 | IZO | 69.7943 | 39.5377 | 129.2194 | 284.6831 | 282.8247 | 284.5298 | 286.1423 | 284.9288 | 0.45436 | 0.23347 | 0.45765 | 0.10114 | 0.46093 | 0.50816 | 2016-09-14 08:56 | 2.2009 | 2.2227 | 2016-09-14 08:57 | 2.1851 | 2.2063 | 2016-09-14 08:58 | 2.1695 | 2.1902 | 2016-09-14 08:56 | 2.1929 | 2.2145 | 2016-09-14 08:58 | 2.1773 | 2.1982
72.3 | 40.5 | 137.9 | 14 | 9 | 0 | 0 | 9 | 16 | C | 83 | IZO | 67.859 | 38.2663 | 125.5978 | 284.8858 | 280.714 | 284.3976 | 288.1307 | 285.2929 | 0.46748 | 0.23363 | 0.47074 | 0.10039 | 0.47399 | 0.50792 | 2016-09-14 09:00 | 2.1392 | 2.1588 | 2016-09-14 09:01 | 2.1243 | 2.1435 | 2016-09-14 09:02 | 2.1098 | 2.1284 | 2016-09-14 09:00 | 2.1317 | 2.1511 | 2016-09-14 09:02 | 2.117 | 2.1359
70.5 | 39.3 | 134.4 | 14 | 9 | 4 | 0 | 9 | 16 | C | 83 | IZO | 66.2012 | 37.0868 | 122.2582 | 285.9943 | 278.7806 | 284.4203 | 291.5887 | 285.8062 | 0.48047 | 0.23454 | 0.4837 | 0.099697 | 0.48692 | 0.50796 | 2016-09-14 09:04 | 2.0813 | 2.099 | 2016-09-14 09:05 | 2.0674 | 2.0846 | 2016-09-14 09:06 | 2.0537 | 2.0705 | 2016-09-14 09:05 | 2.0743 | 2.0918 | 2016-09-14 09:06 | 2.0605 | 2.0776
68.4 | 38.5 | 131 | 14 | 9 | 8 | 0 | 9 | 16 | C | 83 | IZO | 64.2682 | 36.2973 | 119.0079 | 284.8195 | 281.2684 | 284.1232 | 287.5793 | 284.8127 | 0.49334 | 0.23358 | 0.49654 | 0.10059 | 0.49973 | 0.50743 | 2016-09-14 09:08 | 2.027 | 2.043 | 2016-09-14 09:09 | 2.0139 | 2.0295 | 2016-09-14 09:10 | 2.0011 | 2.0163 | 2016-09-14 09:08 | 2.0204 | 2.0362 | 2016-09-14 09:10 | 2.0075 | 2.0229
66.7 | 37.4 | 128 | 14 | 9 | 12 | 0 | 9 | 16 | C | 83 | IZO | 62.7019 | 35.2077 | 116.1461 | 285.1991 | 278.7776 | 284.3876 | 290.1764 | 285.7676 | 0.50609 | 0.23389 | 0.50925 | 0.099696 | 0.51241 | 0.5079 | 2016-09-14 09:12 | 1.9759 | 1.9904 | 2016-09-14 09:13 | 1.9637 | 1.9778 | 2016-09-14 09:14 | 1.9516 | 1.9653 | 2016-09-14 09:12 | 1.9698 | 1.9841 | 2016-09-14 09:14 | 1.9576 | 1.9715
55.4 | 31.7 | 107.9 | 14 | 9 | 45 | 0 | 9 | 16 | C | 83 | IZO | 52.2135 | 29.5062 | 97.2316 | 284.2579 | 280.2652 | 284.1908 | 287.346 | 285.1526 | 0.60576 | 0.23312 | 0.60861 | 0.10023 | 0.61145 | 0.50755 | 2016-09-14 09:45 | 1.6508 | 1.6576 | 2016-09-14 09:46 | 1.6431 | 1.6497 | 2016-09-14 09:47 | 1.6355 | 1.6419 | 2016-09-14 09:46 | 1.6469 | 1.6536 | 2016-09-14 09:47 | 1.6393 | 1.6458
54.5 | 30.8 | 105.3 | 14 | 9 | 50 | 0 | 9 | 16 | C | 83 | IZO | 51.3708 | 28.5985 | 94.8133 | 286.9364 | 276.17 | 283.3863 | 295.261 | 285.1724 | 0.61992 | 0.23531 | 0.62271 | 0.098764 | 0.6255 | 0.50611 | 2016-09-14 09:50 | 1.6131 | 1.6192 | 2016-09-14 09:51 | 1.6059 | 1.6118 | 2016-09-14 09:52 | 1.5987 | 1.6045 | 2016-09-14 09:50 | 1.6095 | 1.6155 | 2016-09-14 09:52 | 1.6023 | 1.6081
52.9 | 30.2 | 102.9 | 14 | 9 | 55 | 0 | 9 | 16 | C | 83 | IZO | 49.8674 | 27.992 | 92.5697 | 284.0473 | 276.5031 | 282.6835 | 289.8796 | 284.2105 | 0.6338 | 0.23294 | 0.63655 | 0.098883 | 0.63928 | 0.50486 | 2016-09-14 09:55 | 1.5778 | 1.5832 | 2016-09-14 09:56 | 1.571 | 1.5763 | 2016-09-14 09:57 | 1.5643 | 1.5694 | 2016-09-14 09:56 | 1.5744 | 1.5797 | 2016-09-14 09:56 | 1.5676 | 1.5729
51.7 | 30 | 101.6 | 14 | 9 | 59 | 0 | 9 | 16 | C | 83 | IZO | 48.7352 | 27.7896 | 91.3495 | 281.8029 | 281.3922 | 283.8602 | 282.1187 | 284.458 | 0.64471 | 0.2311 | 0.64741 | 0.10063 | 0.6501 | 0.50696 | 2016-09-14 09:59 | 1.5511 | 1.556 | 2016-09-14 10:00 | 1.5446 | 1.5495 | 2016-09-14 10:01 | 1.5382 | 1.543 | 2016-09-14 09:59 | 1.5478 | 1.5528 | 2016-09-14 10:00 | 1.5414 | 1.5462
51.1 | 29.4 | 99.9 | 14 | 10 | 3 | 0 | 9 | 16 | C | 83 | IZO | 48.1681 | 27.1815 | 89.7511 | 283.6676 | 278.5496 | 283.3966 | 287.6227 | 284.5903 | 0.65544 | 0.23263 | 0.65809 | 0.099615 | 0.66073 | 0.50613 | 2016-09-14 10:03 | 1.5257 | 1.5303 | 2016-09-14 10:04 | 1.5195 | 1.524 | 2016-09-14 10:05 | 1.5135 | 1.5178 | 2016-09-14 10:04 | 1.5226 | 1.5271 | 2016-09-14 10:04 | 1.5165 | 1.5209
50.2 | 29.1 | 98.6 | 14 | 10 | 7 | 0 | 9 | 16 | C | 83 | IZO | 47.3161 | 26.877 | 88.528 | 282.9607 | 280.8776 | 284.0603 | 284.5683 | 284.8374 | 0.66598 | 0.23205 | 0.66859 | 0.10045 | 0.67118 | 0.50732 | 2016-09-14 10:07 | 1.5015 | 1.5057 | 2016-09-14 10:08 | 1.4957 | 1.4998 | 2016-09-14 10:09 | 1.4899 | 1.4939 | 2016-09-14 10:08 | 1.4986 | 1.5028 | 2016-09-14 10:08 | 1.4928 | 1.4968
49.5 | 28.7 | 97 | 14 | 10 | 11 | 0 | 9 | 16 | C | 83 | IZO | 46.652 | 26.4706 | 87.0233 | 283.4759 | 280.9467 | 283.4124 | 285.428 | 284.0102 | 0.67633 | 0.23248 | 0.67889 | 0.10047 | 0.68144 | 0.50616 | 2016-09-14 10:11 | 1.4786 | 1.4824 | 2016-09-14 10:12 | 1.473 | 1.4768 | 2016-09-14 10:13 | 1.4675 | 1.4712 | 2016-09-14 10:12 | 1.4758 | 1.4796 | 2016-09-14 10:12 | 1.4702 | 1.474
48.7 | 28.3 | 95.8 | 14 | 10 | 15 | 0 | 9 | 16 | C | 83 | IZO | 45.8911 | 26.0635 | 85.8935 | 282.9049 | 280.6269 | 283.9496 | 284.6623 | 284.7621 | 0.6865 | 0.23201 | 0.68901 | 0.10036 | 0.6915 | 0.50712 | 2016-09-14 10:15 | 1.4567 | 1.4602 | 2016-09-14 10:16 | 1.4514 | 1.4548 | 2016-09-14 10:17 | 1.4461 | 1.4495 | 2016-09-14 10:16 | 1.454 | 1.4575 | 2016-09-14 10:16 | 1.4487 | 1.4522
48.2 | 28 | 94.4 | 14 | 10 | 19 | 0 | 9 | 16 | C | 83 | IZO | 45.4145 | 25.7578 | 84.5743 | 284.45 | 281.9146 | 283.5216 | 286.4059 | 283.9048 | 0.69646 | 0.23327 | 0.69892 | 0.10082 | 0.70137 | 0.50636 | 2016-09-14 10:19 | 1.4358 | 1.4391 | 2016-09-14 10:20 | 1.4308 | 1.434 | 2016-09-14 10:21 | 1.4258 | 1.4289 | 2016-09-14 10:20 | 1.4333 | 1.4365 | 2016-09-14 10:20 | 1.4283 | 1.4314
47.4 | 27.6 | 93.2 | 14 | 10 | 23 | 0 | 9 | 16 | C | 83 | IZO | 44.6498 | 25.3496 | 83.4445 | 283.289 | 280.8585 | 283.6081 | 285.1635 | 284.2774 | 0.70623 | 0.23232 | 0.70864 | 0.10044 | 0.71103 | 0.50651 | 2016-09-14 10:23 | 1.416 | 1.419 | 2016-09-14 10:24 | 1.4112 | 1.4141 | 2016-09-14 10:25 | 1.4064 | 1.4093 | 2016-09-14 10:24 | 1.4136 | 1.4165 | 2016-09-14 10:24 | 1.4088 | 1.4117
47.3 | 27.2 | 92.2 | 14 | 10 | 27 | 0 | 9 | 16 | C | 83 | IZO | 44.554 | 24.9409 | 82.5048 | 287.675 | 279.4253 | 284.2451 | 294.0486 | 285.4328 | 0.71579 | 0.23592 | 0.71815 | 0.099928 | 0.7205 | 0.50765 | 2016-09-14 10:27 | 1.3971 | 1.3998 | 2016-09-14 10:28 | 1.3925 | 1.3952 | 2016-09-14 10:29 | 1.3879 | 1.3906 | 2016-09-14 10:28 | 1.3948 | 1.3975 | 2016-09-14 10:28 | 1.3902 | 1.3928
46.3 | 26.8 | 90.9 | 14 | 10 | 31 | 0 | 9 | 16 | C | 83 | IZO | 43.5939 | 24.5316 | 81.2875 | 284.2831 | 277.6198 | 283.5529 | 289.4297 | 285.0195 | 0.72515 | 0.23314 | 0.72746 | 0.099282 | 0.72975 | 0.50641 | 2016-09-14 10:31 | 1.379 | 1.3815 | 2016-09-14 10:32 | 1.3747 | 1.3771 | 2016-09-14 10:33 | 1.3703 | 1.3727 | 2016-09-14 10:32 | 1.3768 | 1.3793 | 2016-09-14 10:32 | 1.3725 | 1.3749
45.6 | 26.6 | 89.8 | 14 | 10 | 35 | 0 | 9 | 16 | C | 83 | IZO | 42.9193 | 24.3267 | 80.2611 | 283.1197 | 279.6692 | 283.4299 | 285.7818 | 284.3527 | 0.7343 | 0.23218 | 0.73656 | 0.10001 | 0.7388 | 0.50619 | 2016-09-14 10:35 | 1.3618 | 1.3642 | 2016-09-14 10:36 | 1.3577 | 1.3599 | 2016-09-14 10:37 | 1.3536 | 1.3558 | 2016-09-14 10:36 | 1.3597 | 1.362 | 2016-09-14 10:36 | 1.3556 | 1.3578
45.3 | 26.4 | 88.8 | 14 | 10 | 39 | 0 | 9 | 16 | C | 83 | IZO | 42.6296 | 24.1217 | 79.3302 | 285.1855 | 281.4668 | 283.5079 | 288.0546 | 284.0003 | 0.74324 | 0.23388 | 0.74544 | 0.10066 | 0.74763 | 0.50633 | 2016-09-14 10:39 | 1.3455 | 1.3476 | 2016-09-14 10:40 | 1.3415 | 1.3436 | 2016-09-14 10:41 | 1.3376 | 1.3396 | 2016-09-14 10:40 | 1.3435 | 1.3456 | 2016-09-14 10:40 | 1.3395 | 1.3416
44.7 | 26.2 | 87.9 | 14 | 10 | 43 | 0 | 9 | 16 | C | 83 | IZO | 42.0495 | 23.9165 | 78.4938 | 284.413 | 283.0144 | 283.7992 | 285.4877 | 283.9771 | 0.75197 | 0.23324 | 0.75411 | 0.10121 | 0.75625 | 0.50685 | 2016-09-14 10:43 | 1.3298 | 1.3318 | 2016-09-14 10:44 | 1.3261 | 1.328 | 2016-09-14 10:45 | 1.3223 | 1.3242 | 2016-09-14 10:44 | 1.3279 | 1.3299 | 2016-09-14 10:44 | 1.3242 | 1.3261
44.2 | 25.8 | 86.9 | 14 | 10 | 47 | 0 | 9 | 16 | C | 83 | IZO | 41.5651 | 23.5059 | 77.5658 | 284.2941 | 279.9338 | 283.5221 | 287.6587 | 284.402 | 0.76047 | 0.23315 | 0.76257 | 0.10011 | 0.76464 | 0.50636 | 2016-09-14 10:47 | 1.315 | 1.3168 | 2016-09-14 10:48 | 1.3114 | 1.3131 | 2016-09-14 10:49 | 1.3078 | 1.3095 | 2016-09-14 10:48 | 1.3132 | 1.3149 | 2016-09-14 10:48 | 1.3096 | 1.3113
43.9 | 25.5 | 86.2 | 14 | 10 | 51 | 0 | 9 | 16 | C | 83 | IZO | 41.274 | 23.1975 | 76.9167 | 285.7763 | 278.7207 | 284.2699 | 291.2247 | 285.6409 | 0.76876 | 0.23436 | 0.7708 | 0.099676 | 0.77282 | 0.50769 | 2016-09-14 10:51 | 1.3008 | 1.3024 | 2016-09-14 10:52 | 1.2974 | 1.299 | 2016-09-14 10:53 | 1.294 | 1.2955 | 2016-09-14 10:52 | 1.2991 | 1.3007 | 2016-09-14 10:52 | 1.2957 | 1.2972
43.4 | 25.3 | 85.4 | 14 | 10 | 55 | 0 | 9 | 16 | C | 83 | IZO | 40.7881 | 22.9918 | 76.1754 | 285.2424 | 279.4566 | 284.459 | 289.7088 | 285.6931 | 0.77683 | 0.23392 | 0.77881 | 0.099939 | 0.78078 | 0.50803 | 2016-09-14 10:55 | 1.2873 | 1.2888 | 2016-09-14 10:56 | 1.284 | 1.2855 | 2016-09-14 10:57 | 1.2808 | 1.2822 | 2016-09-14 10:56 | 1.2856 | 1.2871 | 2016-09-14 10:56 | 1.2824 | 1.2838
42.9 | 25.2 | 84.8 | 14 | 10 | 59 | 0 | 9 | 16 | C | 83 | IZO | 40.3012 | 22.8888 | 75.6196 | 284.4924 | 282.2169 | 285.3053 | 286.2443 | 286.0603 | 0.78467 | 0.23331 | 0.78659 | 0.10093 | 0.7885 | 0.50954 | 2016-09-14 10:59 | 1.2744 | 1.2758 | 2016-09-14 11:00 | 1.2713 | 1.2726 | 2016-09-14 11:01 | 1.2682 | 1.2695 | 2016-09-14 11:00 | 1.2729 | 1.2742 | 2016-09-14 11:01 | 1.2698 | 1.2711
42.9 | 25.7 | 84.1 | 14 | 11 | 3 | 0 | 9 | 16 | C | 83 | IZO | 40.3012 | 23.4031 | 74.9718 | 288.2424 | 298.4916 | 285.5971 | 280.3091 | 282.3507 | 0.79228 | 0.23638 | 0.79415 | 0.10675 | 0.796 | 0.51006 | 2016-09-14 11:03 | 1.2622 | 1.2634 | 2016-09-14 11:04 | 1.2592 | 1.2604 | 2016-09-14 11:05 | 1.2563 | 1.2575 | 2016-09-14 11:04 | 1.2607 | 1.2619 | 2016-09-14 11:04 | 1.2577 | 1.2589
42.2 | 25.1 | 83.4 | 14 | 11 | 7 | 0 | 9 | 16 | C | 83 | IZO | 39.6177 | 22.7859 | 74.3252 | 285.2149 | 289.4701 | 285.7432 | 281.9164 | 284.7919 | 0.79967 | 0.2339 | 0.80147 | 0.10352 | 0.80327 | 0.51032 | 2016-09-14 11:07 | 1.2505 | 1.2517 | 2016-09-14 11:08 | 1.2477 | 1.2488 | 2016-09-14 11:09 | 1.2449 | 1.246 | 2016-09-14 11:08 | 1.2491 | 1.2502 | 2016-09-14 11:08 | 1.2463 | 1.2474
41.8 | 24.8 | 82.4 | 14 | 11 | 11 | 0 | 9 | 16 | C | 83 | IZO | 39.2263 | 22.4768 | 73.4026 | 284.8259 | 287.018 | 284.4885 | 283.1224 | 283.8372 | 0.80682 | 0.23358 | 0.80857 | 0.10264 | 0.8103 | 0.50808 | 2016-09-14 11:11 | 1.2394 | 1.2405 | 2016-09-14 11:12 | 1.2368 | 1.2378 | 2016-09-14 11:13 | 1.2341 | 1.2351 | 2016-09-14 11:12 | 1.2381 | 1.2391 | 2016-09-14 11:12 | 1.2354 | 1.2364

## param
-0.9 | 1e-05 |  | -1.9 | 0
0 | 1 |  | -1 | 1
0.9 | 1e-05 |  | 1 | 1
|  |  | 1.9 | 0
304.6 | 1e-07 | 310.55 | 1e-06 | 316.7 | 1e-06 | 323.5 | 1e-06 | 323.1 | 1e-06 | 330.5 | 1e-06 | 338 | 1e-06
305.5 | 1000 | 311.45 | 1000 | 317.6 | 1000 | 324.4 | 1000 | 324 | 1000 | 331.4 | 1000 | 338.9 | 1000
306.4 | 1e-07 | 312.35 | 1e-06 | 318.5 | 1e-06 | 326.4 | 1000 | 326 | 1000 | 333.4 | 1000 | 340.9 | 1000
[6 empty] | 327.3 | 1e-06 | 326.9 | 1e-06 | 334.3 | 1e-06 | 341.8 | 1e-06
1.1111
304.6 | 0.001 | 310.55 | 0.001 | 316.7 | 0.001 | 323.5 | 0 | 323.1 | 0.001 | 330.5 | 0.001 | 338 | 1e-07
305.5 | 100 | 311.45 | 100 | 317.6 | 100 | 324.4 | 100 | 324 | 100 | 331.4 | 100 | 338.9 | 100
306.4 | 0.001 | 312.35 | 0.001 | 318.5 | 0.001 | 326.4 | 100 | 326 | 100 | 333.4 | 100 | 340.9 | 100
[6 empty] | 327.3 | 0 | 326.9 | 0.001 | 334.3 | 0.001 | 341.8 | 1e-07
58 | 0
92 | 0
code | number | name | Dobson | earth_radius | LONG | Hemisphere | degs | min | sec | LAT | Hemisphere | deg | min | sec | TimeZone | stn Pressure | elevation | Ozone_Height | LongRadian | LatRadian
IZO | 999 | Izana Observatory,Spain | 83 | 6373593 | 16.4973 | W | 16 | 29 | 25 | 28.2986 | N | 28 | 17 | 55 | 0 | 770 | 2360 | 26000 | 0.28781 | 0.4939
=IF( C13<90,(earth_radius+Ozone_height)/(SQRT(((earth_radius+Ozone_height)^2)-((earth_radius+Elevation)^2)*(SIN(RADIANS(C13)))^2)),-90000)
Public Function AirMass#(CosZenAng As Double) | [14 empty] | p0 | 0.75993
AirMass# = 1# / (CosZenAng + 0.15 / (90 - RadsToDegs#(ArcCos#(CosZenAng)) + 3.885) ^ 1.253)
End Function
Ozone Absorption and Molecular Scattering coefficients for Use
with Dobson ozone Spectrophotometers Beginning 1 January 1992
Wavelength | ALFA |  | BETA | d(BETA)/d(ALFA) | [5 empty] | IUP Corr-factor | B&BP Corr_factor | D&M |  |  | α_Komhyr r adj
(nm) | (atm-cm)-1 |  | (atm)-1 | (atm-cm/atm) |  |  |  |  | A_ pair | 0.9898 | 0.9969 | 0.9889 |  |  | 1.806
[9 empty] | C_pair | 0.9922 | 0.9983 | 0.9845 |  |  | 0.833
305 | 0.489 | [7 empty] | D_pair | 0.9642 | 0.9825 | 0.9562 |  |  | 0.374
325 | 0.375 | [7 empty] | AD_pair | 0.9965 | 1.0006 | 0.9975 |  |  | 1.432
Alfa_A | 1.786 | Beta_A | 114 | 63 | 1000 |  |  |  | CD_pair | 1.015 | 1.0112 | 1.0076 |  |  | 0.459
|  |  | 0 | 0
308.9 | 0.466 |  | 0 | 0
329.1 | 0.355 |  | 0 | 0
Alfa_B | 1.192 | Beta_B | 111 | 93
|  |  | 0 | 0
311.5 | 0.45 |  | 0 | 0 |  |  | 5
332.4 | 0.341 |  | 0 | 0
Alfa_C | 0.82009 | Beta_C | 109 | 131
|  |  | 0 | 0
317.5 | 0.414 |  | 0 | 0
339.9 | 0.31 |  | 0 | 0
Alfa_D | 0.35762 | Beta_D | 104 | 278
|  |  | 0 | 0
Alfa_AD | 1.4284 | Beta_AD | 10 | 7
Alfa_BD | 0.818 | Beta_BD | 7 | 9
Alfa_CD | 0.46249 | Beta_CD | 5 | 11

## Nfunc083
D83RA | D83NA | D83RC | D83NC | D83RD | D83ND | Rby5 | a | c | d | in pretty print
5 | -1.51 | 5 | 2.18 | 5 | 3.26 | 5 | -1.51 | 2.18 | 3.26 | -1.5 | 2.2 | 3.3 | 0.01 | 0.02 | 0.04
10.74 | 4.7 | 10.76 | 8.39 | 10.72 | 9.62 | 15 | 9.2606 | 12.8971 | 14.1824 | 9.2 | 12.9 | 14.2 | -0.060647 | 0.0029402 | 0.017648
16.48 | 10.83 | 16.51 | 14.49 | 16.44 | 15.7 | 25 | 19.7018 | 23.3471 | 24.6275 | 19.7 | 23.3 | 24.6 | -0.001759 | -0.047139 | -0.027526
22.22 | 16.83 | 22.27 | 20.51 | 22.16 | 21.69 | 35 | 29.9065 | 33.5082 | 34.7616 | 29.9 | 33.5 | 34.8 | -0.0065067 | -0.0081693 | 0.038446
27.96 | 22.74 | 28.02 | 26.46 | 27.88 | 27.58 | 45 | 39.8449 | 43.394 | 44.5867 | 39.8 | 43.4 | 44.6 | -0.044871 | 0.0060056 | 0.013291
33.7 | 28.59 | 33.78 | 32.28 | 33.6 | 33.36 | 55 | 49.3404 | 52.8934 | 54.076 | 49.3 | 52.9 | 54.1 | -0.040402 | 0.0065573 | 0.02402
39.87 | 34.8 | 39.99 | 38.5 | 40.01 | 39.73 | 65 | 58.6907 | 62.1867 | 63.3142 | 58.7 | 62.2 | 63.3 | 0.0093361 | 0.013335 | -0.014152
46.12 | 40.93 | 46.25 | 44.6 | 46.26 | 45.8 | 75 | 67.9965 | 71.4015 | 72.4459 | 68 | 71.4 | 72.4 | 0.0034853 | -0.0015358 | -0.045888
52.43 | 46.93 | 52.58 | 50.62 | 52.56 | 51.79 | 85 | 77.243 | 80.6092 | 81.6241 | 77.2 | 80.6 | 81.6 | -0.043043 | -0.0091746 | -0.024141
58.76 | 52.85 | 58.93 | 56.56 | 58.88 | 57.68 | 95 | 86.5778 | 89.8754 | 90.8361 | 86.6 | 89.9 | 90.8 | 0.022224 | 0.024624 | -0.036083
65.01 | 58.7 | 65.22 | 62.39 | 65.17 | 63.47 | 105 | 95.9718 | 99.1641 | 100.0591 | 96 | 99.2 | 100.1 | 0.028156 | 0.03591 | 0.040888
71.68 | 64.91 | 71.96 | 68.6 | 72.14 | 69.83 | 115 | 105.3164 | 108.4291 | 109.276 | 105.3 | 108.4 | 109.3 | -0.016442 | -0.029128 | 0.024023
78.27 | 71.03 | 78.58 | 74.7 | 78.77 | 75.9 | 125 | 114.7278 | 117.7197 | 118.4987 | 114.7 | 117.7 | 118.5 | -0.027761 | -0.019725 | 0.0013422
84.77 | 77.03 | 85.12 | 80.72 | 85.3 | 81.9 | 135 | 124.2697 | 127.0863 | 127.7279 | 124.3 | 127.1 | 127.7 | 0.030284 | 0.013723 | -0.027928
91.14 | 82.95 | 91.54 | 86.67 | 91.69 | 87.78 | 145 | 133.8086 | 136.5037 | 137.0681 | 133.8 | 136.5 | 137.1 | -0.0086133 | -0.0037106 | 0.031879
97.36 | 88.8 | 97.82 | 92.49 | 97.96 | 93.57 | 155 | 143.4602 | 145.9931 | 146.5049 | 143.5 | 146 | 146.5 | 0.03978 | 0.0069477 | -0.0049043
103.97 | 95.01 | 104.5 | 98.7 | 104.86 | 99.93 | 165 | 153.1836 | 155.5627 | 155.9752 | 153.2 | 155.6 | 156 | 0.016436 | 0.037317 | 0.02477
110.54 | 101.13 | 111.08 | 104.8 | 111.45 | 106.01 | 175 | 162.9608 | 165.2327 | 165.5125 | 163 | 165.2 | 165.5 | 0.039212 | -0.032674 | -0.012494
116.94 | 107.14 | 117.58 | 110.82 | 117.96 | 112 | 185 | 172.7866 | 174.9334 | 175.074 | 172.8 | 174.9 | 175.1 | 0.013394 | -0.033433 | 0.026028
123.23 | 113.05 | 123.98 | 116.77 | 124.34 | 117.89 | 195 | 182.6394 | 184.6057 | 184.6243 | 182.6 | 184.6 | 184.6 | -0.039352 | -0.0057191 | -0.024324
129.38 | 118.9 | 130.22 | 122.59 | 130.61 | 123.67 | 205 | 192.4841 | 194.2836 | 194.1583 | 192.5 | 194.3 | 194.2 | 0.015876 | 0.016413 | 0.041746
135.88 | 125.11 | 136.83 | 128.81 | 137.48 | 130.03 | 215 | 202.4109 | 204.0518 | 203.7711 | 202.4 | 204.1 | 203.8 | -0.010885 | 0.048175 | 0.028851
142.31 | 131.24 | 143.3 | 134.9 | 143.98 | 136.11 | 225 | 212.4107 | 213.9099 | 213.4697 | 212.4 | 213.9 | 213.5 | -0.010729 | -0.0099222 | 0.030315
148.58 | 137.24 | 149.67 | 140.92 | 150.34 | 142.1 | 235 | 222.6125 | 223.9066 | 223.3561 | 222.6 | 223.9 | 223.4 | -0.01248 | -0.0066373 | 0.043946
154.69 | 143.16 | 155.92 | 146.87 | 156.57 | 147.99 | 245 | 232.9094 | 234.043 | 233.3305 | 232.9 | 234 | 233.3 | -0.0094379 | -0.04302 | -0.030472
160.72 | 149 | 162.02 | 152.69 | 162.68 | 153.77 | 255 | 243.0375 | 244.1782 | 243.3484 | 243 | 244.2 | 243.3 | -0.037453 | 0.021839 | -0.048358
167.07 | 155.21 | 168.46 | 158.91 | 169.37 | 160.14 | 265 | 253.2588 | 254.2376 | 253.357 | 253.3 | 254.2 | 253.4 | 0.041246 | -0.037564 | 0.043032
173.34 | 161.34 | 174.77 | 165.01 | 175.73 | 166.21 | 275 | 263.5293 | 264.3569 | 263.3423 | 263.5 | 264.4 | 263.3 | -0.029307 | 0.043099 | -0.042337
179.48 | 167.34 | 180.98 | 171.03 | 181.99 | 172.2 | 285 | 273.6505 | 274.5418 | 273.4295 | 273.7 | 274.5 | 273.4 | 0.049496 | -0.041795 | -0.029511
185.48 | 173.26 | 187.11 | 176.98 | 188.16 | 178.09 |  |  |  |  | 284 | 284.8 | 283.6 | -0.080761 | -0.032459 | 0.10211
191.42 | 179.11 | 193.13 | 182.8 | 194.22 | 183.88
197.72 | 185.32 | 199.56 | 189.01 | 200.89 | 190.24
203.94 | 191.44 | 205.85 | 195.11 | 207.25 | 196.31
210.02 | 197.44 | 212.02 | 201.13 | 213.48 | 202.31
215.95 | 203.36 | 218.08 | 207.08 | 219.6 | 208.2
221.81 | 209.21 | 223.98 | 212.9 | 225.52 | 213.98
227.98 | 215.42 | 230.23 | 219.11 | 231.96 | 220.34
233.97 | 221.55 | 236.29 | 225.21 | 238.08 | 226.42
239.77 | 227.55 | 242.22 | 231.23 | 244.08 | 232.41
245.54 | 233.46 | 248.1 | 237.18 | 249.96 | 238.3
251.31 | 239.31 | 253.83 | 243 | 255.73 | 244.08
257.45 | 245.52 | 260.03 | 249.22 | 262.08 | 250.45
263.44 | 251.65 | 266.07 | 255.32 | 268.18 | 256.52
269.25 | 257.65 | 272.01 | 261.34 | 274.17 | 262.51
275.04 | 263.57 | 277.89 | 267.28 | 280.03 | 268.4
280.82 | 269.42 | 283.59 | 273.11 | 285.75 | 274.19
286.94 | 275.63 | 289.71 | 279.32 | 292.01 | 280.55
292.84 | 281.75 | 295.61 | 285.42 | 298.01 | 286.62
298.49 | 287.75

## comp of all values
Komyhr adjusted |  |  |  |  | IUP multiplied |  |  |  |  | B&P B multiply |  |  |  |  | D&M multiply
282.5771 | 288.7099 | 282.5144 | 277.6237 | 280.8825 | 280.373 | 278.3741 | 279.6328 | 281.788 | 279.8994 | 282.0967 | 283.6575 | 281.6386 | 280.733 | 281.051 | 278.1972 | 276.0644 | 279.3785 | 279.7336 | 280.1803
D083_TO3_C | D083_TO3_D | D083_TO3_A | D083_TO3_CD | D083_TO3_AD | D083_TO3_C | D083_TO3_D | D083_TO3_A | D083_TO3_CD | D083_TO3_AD | D083_TO3_C | D083_TO3_D | D083_TO3_A | D083_TO3_CD | D083_TO3_AD | D083_TO3_C | D083_TO3_D | D083_TO3_A | D083_TO3_CD | D083_TO3_AD
281.1643 | 283.8618 | 280.6383 | 279.2087 | 279.7205 | 283.3746 | 294.4013 | 283.5303 | 275.0825 | 280.7029 | 281.6431 | 288.9178 | 281.511 | 276.1162 | 279.5527 | 285.5909 | 296.8644 | 283.7883 | 277.1027 | 280.4215
281.0826 | 284.8637 | 281.1767 | 278.2046 | 280.15 | 283.2923 | 295.4405 | 284.0743 | 274.0932 | 281.134 | 281.5612 | 289.9376 | 282.0511 | 275.1232 | 279.982 | 285.5079 | 297.9123 | 284.3328 | 276.1062 | 280.8522
281.5811 | 282.884 | 281.5984 | 280.6911 | 281.2086 | 283.7947 | 293.3872 | 284.5003 | 276.5429 | 282.1963 | 282.0606 | 287.9226 | 282.474 | 277.5822 | 281.04 | 286.0143 | 295.8418 | 284.7592 | 278.5739 | 281.9134
281.0019 | 283.6208 | 281.3255 | 279.0142 | 280.6799 | 283.2109 | 294.1514 | 284.2246 | 274.8909 | 281.6657 | 281.4804 | 288.6725 | 282.2003 | 275.9239 | 280.5116 | 285.426 | 296.6124 | 284.4832 | 276.9097 | 281.3833
280.5029 | 283.2087 | 281.1044 | 278.4239 | 280.5151 | 282.708 | 293.724 | 284.0012 | 274.3093 | 281.5004 | 280.9806 | 288.2531 | 281.9785 | 275.3401 | 280.3469 | 284.9192 | 296.1814 | 284.2597 | 276.3238 | 281.2182
280.9828 | 283.4613 | 281.3154 | 279.072 | 280.7206 | 283.1917 | 293.986 | 284.2144 | 274.9478 | 281.7066 | 281.4613 | 288.5102 | 282.1902 | 275.981 | 280.5523 | 285.4066 | 296.4456 | 284.4731 | 276.9671 | 281.4241
281.3114 | 283.1242 | 281.3012 | 279.929 | 280.795 | 283.5229 | 293.6364 | 284.2 | 275.7921 | 281.7812 | 281.7905 | 288.1672 | 282.1759 | 276.8286 | 280.6266 | 285.7404 | 296.0931 | 284.4586 | 277.8176 | 281.4987
281.6817 | 284.1741 | 282.1693 | 279.7338 | 281.6194 | 283.8961 | 294.7253 | 285.0771 | 275.5998 | 282.6085 | 282.1614 | 289.2358 | 283.0468 | 276.6355 | 281.4505 | 286.1165 | 297.1911 | 285.3366 | 277.6239 | 282.3252
281.5476 | 286.1065 | 282.2343 | 277.9061 | 281.1997 | 283.761 | 296.7294 | 285.1427 | 273.7991 | 282.1874 | 282.0271 | 291.2025 | 283.1119 | 274.828 | 281.0311 | 285.9803 | 299.212 | 285.4022 | 275.81 | 281.9045
281.8017 | 285.4359 | 281.7318 | 278.9051 | 280.7438 | 284.017 | 296.0339 | 284.6351 | 274.7834 | 281.7299 | 282.2815 | 290.52 | 282.6079 | 275.816 | 280.5755 | 286.2384 | 298.5107 | 284.8942 | 276.8014 | 281.4475
282.1712 | 285.0323 | 281.9098 | 279.8974 | 281.076 | 284.3894 | 295.6154 | 284.815 | 275.761 | 282.0632 | 282.6517 | 290.1092 | 282.7865 | 276.7973 | 280.9075 | 286.6137 | 298.0886 | 285.0742 | 277.7863 | 281.7804
281.5122 | 286.1608 | 282.3418 | 277.7759 | 281.328 | 283.7253 | 296.7857 | 285.2514 | 273.6709 | 282.3161 | 281.9916 | 291.2578 | 283.2198 | 274.6993 | 281.1593 | 285.9444 | 299.2688 | 285.511 | 275.6808 | 282.0331
282.6118 | 285.9516 | 282.503 | 279.9366 | 281.5876 | 284.8335 | 296.5687 | 285.4143 | 275.7996 | 282.5767 | 283.0931 | 291.0449 | 283.3815 | 276.8361 | 281.4188 | 287.0613 | 299.05 | 285.674 | 277.8251 | 282.2934
282.3819 | 286.6367 | 282.3341 | 278.9565 | 281.1971 | 284.6018 | 297.2793 | 285.2436 | 274.834 | 282.1847 | 282.8627 | 291.7422 | 283.212 | 275.8668 | 281.0285 | 286.8277 | 299.7665 | 285.5032 | 276.8524 | 281.9018
282.6981 | 283.8365 | 282.7173 | 281.808 | 282.413 | 284.9205 | 294.3752 | 285.6307 | 277.6433 | 283.4049 | 283.1795 | 288.8921 | 283.5964 | 278.6867 | 282.2436 | 287.1489 | 296.838 | 285.8907 | 279.6824 | 283.1208
282.2028 | 285.667 | 282.745 | 279.4141 | 281.971 | 284.4213 | 296.2736 | 285.6587 | 275.2849 | 282.9614 | 282.6834 | 290.7552 | 283.6243 | 276.3194 | 281.8019 | 286.6458 | 298.7524 | 285.9187 | 277.3066 | 282.6777
282.4581 | 284.0842 | 282.6532 | 281.1641 | 282.2695 | 284.6786 | 294.632 | 285.5659 | 277.009 | 283.261 | 282.9391 | 289.1442 | 283.5321 | 278.0499 | 282.1003 | 286.9052 | 297.097 | 285.8258 | 279.0434 | 282.977
283.6048 | 282.6669 | 282.7142 | 284.3971 | 282.7175 | 285.8343 | 293.1621 | 285.6276 | 280.1942 | 283.7105 | 284.0877 | 287.7017 | 283.5933 | 281.2471 | 282.548 | 288.0699 | 295.6148 | 285.8875 | 282.252 | 283.4261
282.5029 | 285.4731 | 282.4585 | 280.1084 | 281.663 | 284.7237 | 296.0725 | 285.3693 | 275.9689 | 282.6522 | 282.984 | 290.5579 | 283.3369 | 277.0059 | 281.4941 | 286.9506 | 298.5496 | 285.629 | 277.9956 | 282.3689
282.9308 | 283.5145 | 282.7578 | 282.4788 | 282.5526 | 285.1551 | 294.0411 | 285.6717 | 278.3043 | 283.545 | 283.4126 | 288.5643 | 283.6371 | 279.3501 | 282.3832 | 287.3853 | 296.5012 | 285.9316 | 280.3482 | 283.2608
282.4281 | 288.2434 | 282.8583 | 277.7021 | 281.4479 | 284.6484 | 298.9456 | 285.7732 | 273.5981 | 282.4364 | 282.9091 | 293.3775 | 283.7379 | 274.6263 | 281.2791 | 286.8747 | 301.4467 | 286.0333 | 275.6075 | 282.1533
285.1252 | 284.7969 | 282.1046 | 285.4041 | 281.3978 | 287.3667 | 295.3712 | 285.0117 | 281.1863 | 282.3862 | 285.6108 | 289.8696 | 282.9818 | 282.243 | 281.2291 | 289.6143 | 297.8424 | 285.2711 | 283.2514 | 282.1031
282.34 | 285.5757 | 281.4506 | 279.7138 | 280.3699 | 284.5596 | 296.1789 | 284.351 | 275.5801 | 281.3547 | 282.8208 | 290.6623 | 282.3258 | 276.6157 | 280.2018 | 286.7852 | 298.6569 | 284.6098 | 277.604 | 281.0726
280.1769 | 290.6122 | 282.6465 | 271.6836 | 280.563 | 282.3794 | 301.4025 | 285.5593 | 267.6686 | 281.5484 | 280.654 | 295.7885 | 283.5255 | 268.6745 | 280.3948 | 284.588 | 303.9241 | 285.8191 | 269.6344 | 281.2662
282.0583 | 288.2496 | 282.2198 | 277.0225 | 280.642 | 284.2756 | 298.9521 | 285.1281 | 272.9286 | 281.6277 | 282.5386 | 293.3838 | 283.0974 | 273.9542 | 280.4738 | 286.499 | 301.4533 | 285.3876 | 274.933 | 281.3454
281.4071 | 290.8249 | 282.9072 | 273.7417 | 280.8366 | 283.6193 | 301.623 | 285.8226 | 269.6963 | 281.823 | 281.8863 | 296.005 | 283.7869 | 270.7098 | 280.6682 | 285.8376 | 304.1465 | 286.0827 | 271.677 | 281.5404
281.9584 | 291.234 | 282.2971 | 274.4084 | 279.9605 | 284.175 | 302.0473 | 285.2062 | 270.3531 | 280.9438 | 282.4385 | 296.4214 | 283.175 | 271.3691 | 279.7926 | 286.3976 | 304.5743 | 285.4658 | 272.3387 | 280.6621
281.4394 | 291.2647 | 282.8583 | 273.4411 | 280.6604 | 283.6519 | 302.0792 | 285.7732 | 269.4001 | 281.6462 | 281.9187 | 296.4526 | 283.7379 | 270.4125 | 280.4921 | 285.8704 | 304.6065 | 286.0333 | 271.3786 | 281.3638
283.003 | 292.826 | 282.4645 | 275.006 | 279.7561 | 285.2278 | 303.6984 | 285.3753 | 270.9419 | 280.7387 | 283.4849 | 298.0417 | 283.3429 | 271.9601 | 279.5884 | 287.4586 | 306.2393 | 285.6351 | 272.9317 | 280.4573
281.9015 | 292.1394 | 282.5788 | 273.5661 | 280.0797 | 284.1176 | 302.9863 | 285.4908 | 269.5232 | 281.0634 | 282.3816 | 297.3429 | 283.4575 | 270.5361 | 279.9117 | 286.3398 | 305.5213 | 285.7506 | 271.5026 | 280.7816
286.2602 | 291.0856 | 283.2369 | 282.3346 | 281.1851 | 288.5106 | 301.8934 | 286.1557 | 278.1622 | 282.1727 | 286.7477 | 296.2703 | 284.1177 | 279.2075 | 281.0164 | 290.7671 | 304.4191 | 286.4161 | 280.205 | 281.8898
282.9606 | 289.6689 | 282.58 | 277.5005 | 280.7267 | 285.1851 | 300.4241 | 285.492 | 273.3995 | 281.7127 | 283.4425 | 294.8284 | 283.4587 | 274.4269 | 280.5584 | 287.4156 | 302.9375 | 285.7519 | 275.4074 | 281.4303
281.8542 | 291.9313 | 282.4853 | 273.6487 | 280.0165 | 284.07 | 302.7705 | 285.3964 | 269.6046 | 281 | 282.3342 | 297.1311 | 283.3638 | 270.6178 | 279.8486 | 286.2918 | 305.3036 | 285.6561 | 271.5847 | 280.7183
283.926 | 293.9459 | 282.5888 | 275.7668 | 279.6209 | 286.1581 | 304.8599 | 285.5009 | 271.6914 | 280.603 | 284.4095 | 299.1816 | 283.4675 | 272.7124 | 279.4533 | 288.3962 | 307.4105 | 285.7607 | 273.6868 | 280.3217
283.2026 | 295.7143 | 282.9025 | 273.0127 | 279.5549 | 285.4289 | 306.6939 | 285.8178 | 268.978 | 280.5367 | 283.6848 | 300.9815 | 283.7822 | 269.9888 | 279.3872 | 287.6613 | 309.2599 | 286.078 | 270.9534 | 280.2555
283.1217 | 293.0499 | 282.6535 | 275.0366 | 279.9368 | 285.3474 | 303.9306 | 285.5663 | 270.972 | 280.92 | 283.6038 | 298.2696 | 283.5324 | 271.9903 | 279.7689 | 287.5792 | 306.4735 | 285.8262 | 272.9621 | 280.6384
284.6162 | 292.1639 | 283.4174 | 278.4705 | 281.1317 | 286.8537 | 303.0118 | 286.3381 | 274.3551 | 282.1191 | 285.1009 | 297.3679 | 284.2987 | 275.3861 | 280.9631 | 289.0972 | 305.5469 | 286.5987 | 276.37 | 281.8363
284.1249 | 293.1341 | 283.6281 | 276.7881 | 281.1441 | 286.3585 | 304.018 | 286.5509 | 272.6976 | 282.1315 | 284.6087 | 298.3554 | 284.5101 | 273.7224 | 280.9755 | 288.5982 | 306.5615 | 286.8117 | 274.7004 | 281.8487
283.4199 | 296.0326 | 284.488 | 273.1467 | 281.4717 | 285.6479 | 307.024 | 287.4197 | 269.11 | 282.4603 | 283.9025 | 301.3054 | 285.3727 | 270.1213 | 281.3029 | 287.8821 | 309.5927 | 287.6813 | 271.0864 | 282.1771
287.1441 | 298.7399 | 284.799 | 277.6992 | 281.1568 | 289.4014 | 309.8318 | 287.7338 | 273.5953 | 282.1443 | 287.633 | 304.0609 | 285.6846 | 274.6234 | 280.9882 | 291.6649 | 312.424 | 287.9957 | 275.6046 | 281.8615
284.1949 | 303.4633 | 284.965 | 268.4981 | 280.1327 | 286.4291 | 314.7307 | 287.9016 | 264.5301 | 281.1166 | 284.6789 | 308.8685 | 285.8512 | 265.5242 | 279.9647 | 288.6693 | 317.3639 | 288.1636 | 266.4729 | 280.8347
283.8423 | 301.3547 | 283.7453 | 269.5762 | 279.1451 | 286.0737 | 312.5438 | 286.6693 | 265.5923 | 280.1256 | 284.3257 | 306.7223 | 284.6276 | 266.5903 | 278.9778 | 288.3112 | 315.1587 | 286.9302 | 267.5428 | 279.8447