## summaryO3 18 Jun 16
Averages of total ozone:
With the exisiting Calibration, Lamp tests included
Average is over the complete Mu range
D083_TO3_C | D083_TO3_D | D083_TO3_A | D083_TO3_CD | D083_TO3_AD
273.241 | 291.7356 | 273.5109 | 258.9778 | 268.8859
4.355 | 22.8496 | 4.6249 | -9.9081
With Langley Plot Calibration
Langley Plot Adjustments and Lamp Test results included
D083_TO3_C* | D083_TO3_D* | D083_TO3_A* | D083_TO3_CD* | D083_TO3_AD*
268.9581 | 272.5002 | 267.3461 | 266.2646 | 266.0256
2.9325 | 6.4746 | 1.3205 | 0.23903

## IdD083
Dobson 083 at IZO 2016.Sep.19 Langley Plot Results
SLC083Na | -0.8 | SLC083Nc | -0.7 | [9 empty] | SLC083Nd | -0.6 | [11 empty] | Correction to
D083 LP adj. A:-1.63 C:-0.71 D:-1.11 Nad:-0.51 | [12 empty] | MaxMu | 2.5 | MaxMu can be changed | [5 empty] | LPC_C_83 |  | LPC_D_83 |  | LPC_A_83 | AD | CD
[20 empty] | With out High mu values | -0.65326 |  | -1.1933 |  | -1.9744 | -0.7811 | 0.54003 | [14 empty] | Limited Mu range
[11 empty] | 46 |  |  |  | 273.241 | 291.7356 | 273.5109 | 258.9778 | 268.8859 | all Values | -0.71138 |  | -1.1118 |  | -1.6266 | -0.51474 | 0.40043 | [14 empty] | 0.53598 |  | -0.77447
D083RC | D083RD | D083RA | day | h | m | s | m | year |  |  |  | D083NC | D083ND | D083NA | D083_TO3_C | D083_TO3_D | D083_TO3_A | D083_TO3_CD | D083_TO3_AD | 1/mu.C | P*C_83 | 1/Mu.D | P*D_83 | 1/Mu.A | P*A_83 | lst.C | Mu.C | AirM.C | lst.D | Mu.D | AirM.D | lst.A | Mu.A | AirM.A | lst.CD | Mu.CD | AirM.CD | lst.AD | Mu.AD | AirM.AD | 1/mu.cd | P*CD_83 | 1/mu.ad | P*AD_83
138 | 79.4 | 245.8 | 19 | 7 | 56 | 0 | 9 | 16 | H | 83 | IZO | 129.2108 | 75.878 | 232.9249 | 270.1557 | 283.0251 | 271.0605 | 260.4646 | 267.9483 | 0.23991 | 0.22328 | 0.24338 | 0.10206 | 0.24685 | 0.48454 | 2016-09-19 07:56 | 4.1682 | 4.3631 | 2016-09-19 07:57 | 4.1089 | 4.2947 | 2016-09-19 07:58 | 4.0511 | 4.2284 | 2016-09-19 07:56 | 4.1383 | 4.3286 | 2016-09-19 07:57 | 4.0798 | 4.2613 | 0.24164 | 0.12122 | 0.24511 | 0.38248
130.3 | 74.9 | 233.7 | 19 | 8 | 0 | 0 | 9 | 16 | H | 83 | IZO | 121.9649 | 71.7544 | 220.4722 | 270.1826 | 284.0369 | 271.2636 | 259.6845 | 267.9611 | 0.25379 | 0.22331 | 0.25727 | 0.10243 | 0.26075 | 0.48491 | 2016-09-19 08:00 | 3.9402 | 4.1019 | 2016-09-19 08:01 | 3.887 | 4.0415 | 2016-09-19 08:02 | 3.8351 | 3.9829 | 2016-09-19 08:01 | 3.9134 | 4.0715 | 2016-09-19 08:02 | 3.8609 | 4.012 | 0.25553 | 0.12088 | 0.25901 | 0.38248
123.5 | 70.6 | 222.9 | 19 | 8 | 4 | 0 | 9 | 16 | H | 83 | IZO | 115.6231 | 67.8244 | 209.5026 | 270.6551 | 283.186 | 271.7945 | 261.1462 | 268.8523 | 0.26771 | 0.2237 | 0.27119 | 0.10212 | 0.27467 | 0.48585 | 2016-09-19 08:04 | 3.7354 | 3.8708 | 2016-09-19 08:05 | 3.6875 | 3.8172 | 2016-09-19 08:06 | 3.6407 | 3.765 | 2016-09-19 08:04 | 3.7113 | 3.8438 | 2016-09-19 08:06 | 3.6639 | 3.7909 | 0.26945 | 0.12158 | 0.27293 | 0.38373
117.1 | 66.9 | 212.5 | 19 | 8 | 8 | 0 | 9 | 16 | H | 83 | IZO | 109.6747 | 64.4522 | 199.1111 | 270.2755 | 283.5706 | 271.5221 | 260.1468 | 268.4228 | 0.28164 | 0.22338 | 0.28512 | 0.10226 | 0.2886 | 0.48537 | 2016-09-19 08:08 | 3.5507 | 3.6651 | 2016-09-19 08:09 | 3.5073 | 3.6171 | 2016-09-19 08:10 | 3.465 | 3.5704 | 2016-09-19 08:08 | 3.5289 | 3.6409 | 2016-09-19 08:10 | 3.486 | 3.5936 | 0.28338 | 0.12112 | 0.28686 | 0.38311
111.4 | 63.6 | 203.2 | 19 | 8 | 12 | 0 | 9 | 16 | H | 83 | IZO | 104.3965 | 61.4298 | 189.9111 | 270.2221 | 284.1127 | 271.6082 | 259.6096 | 268.4009 | 0.29556 | 0.22334 | 0.29904 | 0.10245 | 0.30252 | 0.48552 | 2016-09-19 08:12 | 3.3834 | 3.4808 | 2016-09-19 08:13 | 3.344 | 3.4377 | 2016-09-19 08:14 | 3.3056 | 3.3957 | 2016-09-19 08:12 | 3.3636 | 3.4591 | 2016-09-19 08:13 | 3.3247 | 3.4165 | 0.2973 | 0.12089 | 0.30078 | 0.38307
106.6 | 60.7 | 194.9 | 19 | 8 | 16 | 0 | 9 | 16 | H | 83 | IZO | 99.9483 | 58.762 | 181.7408 | 271.4266 | 285.26 | 272.0371 | 260.8396 | 268.6542 | 0.30947 | 0.22433 | 0.31294 | 0.10287 | 0.31641 | 0.48629 | 2016-09-19 08:16 | 3.2314 | 3.3148 | 2016-09-19 08:17 | 3.1955 | 3.2759 | 2016-09-19 08:18 | 3.1605 | 3.2379 | 2016-09-19 08:17 | 3.2133 | 3.2953 | 2016-09-19 08:18 | 3.1779 | 3.2568 | 0.3112 | 0.12147 | 0.31467 | 0.38342
102 | 58 | 187.1 | 19 | 8 | 20 | 0 | 9 | 16 | H | 83 | IZO | 95.6763 | 56.2646 | 174.0561 | 271.7467 | 285.7577 | 272.0391 | 261.0068 | 268.5359 | 0.32334 | 0.2246 | 0.3268 | 0.10305 | 0.33026 | 0.48629 | 2016-09-19 08:20 | 3.0927 | 3.1647 | 2016-09-19 08:21 | 3.0599 | 3.1294 | 2016-09-19 08:22 | 3.0279 | 3.0949 | 2016-09-19 08:20 | 3.0762 | 3.147 | 2016-09-19 08:22 | 3.0438 | 3.1121 | 0.32507 | 0.12155 | 0.32853 | 0.38324
97.8 | 55.7 | 180.3 | 19 | 8 | 24 | 0 | 9 | 16 | H | 83 | IZO | 91.7714 | 54.1284 | 167.3463 | 272.0491 | 287.6085 | 272.6664 | 260.0979 | 268.8581 | 0.33717 | 0.22485 | 0.34062 | 0.10371 | 0.34407 | 0.48741 | 2016-09-19 08:24 | 2.9659 | 3.0283 | 2016-09-19 08:25 | 2.9358 | 2.9962 | 2016-09-19 08:26 | 2.9064 | 2.9647 | 2016-09-19 08:24 | 2.9508 | 3.0122 | 2016-09-19 08:26 | 2.921 | 2.9804 | 0.3389 | 0.12113 | 0.34234 | 0.3837
93.7 | 53.4 | 173.6 | 19 | 8 | 28 | 0 | 9 | 16 | H | 83 | IZO | 87.9711 | 51.9794 | 160.7941 | 271.4037 | 287.5575 | 272.5079 | 258.9816 | 268.6759 | 0.35095 | 0.22432 | 0.35438 | 0.1037 | 0.35781 | 0.48713 | 2016-09-19 08:28 | 2.8494 | 2.9039 | 2016-09-19 08:29 | 2.8218 | 2.8745 | 2016-09-19 08:30 | 2.7948 | 2.8458 | 2016-09-19 08:28 | 2.8355 | 2.8892 | 2016-09-19 08:29 | 2.8082 | 2.8601 | 0.35267 | 0.12062 | 0.3561 | 0.38343
90.4 | 51.3 | 167.6 | 19 | 8 | 32 | 0 | 9 | 16 | H | 83 | IZO | 84.913 | 50.0016 | 154.9282 | 272.6782 | 287.4755 | 272.6996 | 261.297 | 268.9396 | 0.36466 | 0.22537 | 0.36808 | 0.10367 | 0.37149 | 0.48747 | 2016-09-19 08:32 | 2.7423 | 2.7901 | 2016-09-19 08:33 | 2.7168 | 2.7631 | 2016-09-19 08:34 | 2.6918 | 2.7367 | 2016-09-19 08:33 | 2.7295 | 2.7765 | 2016-09-19 08:34 | 2.7043 | 2.7498 | 0.36637 | 0.1217 | 0.36979 | 0.38381
86.9 | 49.3 | 161.7 | 19 | 8 | 36 | 0 | 9 | 16 | H | 83 | IZO | 81.6668 | 48.1059 | 149.155 | 271.9866 | 286.619 | 272.121 | 260.7253 | 268.4337 | 0.3783 | 0.2248 | 0.3817 | 0.10336 | 0.3851 | 0.48644 | 2016-09-19 08:36 | 2.6434 | 2.6855 | 2016-09-19 08:37 | 2.6198 | 2.6606 | 2016-09-19 08:38 | 2.5968 | 2.6363 | 2016-09-19 08:36 | 2.6316 | 2.673 | 2016-09-19 08:38 | 2.6082 | 2.6484 | 0.38 | 0.12144 | 0.3834 | 0.38308
83.9 | 47.7 | 156.7 | 19 | 8 | 40 | 0 | 9 | 16 | H | 83 | IZO | 78.8945 | 46.5803 | 144.3048 | 272.3778 | 288.2475 | 272.6475 | 260.1522 | 268.6839 | 0.39187 | 0.22512 | 0.39524 | 0.10395 | 0.39862 | 0.48738 | 2016-09-19 08:40 | 2.5519 | 2.5891 | 2016-09-19 08:41 | 2.5301 | 2.5662 | 2016-09-19 08:42 | 2.5087 | 2.5437 | 2016-09-19 08:40 | 2.5409 | 2.5776 | 2016-09-19 08:42 | 2.5193 | 2.5549 | 0.39356 | 0.12118 | 0.39693 | 0.38343
81.1 | 46.1 | 152 | 19 | 8 | 44 | 0 | 9 | 16 | H | 83 | IZO | 76.3182 | 45.0462 | 139.7499 | 272.7255 | 288.4923 | 273.0441 | 260.5739 | 269.1207 | 0.40534 | 0.22541 | 0.4087 | 0.10403 | 0.41204 | 0.48809 | 2016-09-19 08:44 | 2.4671 | 2.5001 | 2016-09-19 08:45 | 2.4468 | 2.4789 | 2016-09-19 08:46 | 2.4269 | 2.4581 | 2016-09-19 08:44 | 2.4569 | 2.4894 | 2016-09-19 08:46 | 2.4368 | 2.4684 | 0.40702 | 0.12137 | 0.41037 | 0.38405
78.4 | 44.6 | 147.4 | 19 | 8 | 48 | 0 | 9 | 16 | H | 83 | IZO | 73.8342 | 43.6005 | 135.3058 | 272.6048 | 288.4858 | 272.9406 | 260.3598 | 268.9943 | 0.41872 | 0.22531 | 0.42205 | 0.10403 | 0.42538 | 0.4879 | 2016-09-19 08:48 | 2.3882 | 2.4176 | 2016-09-19 08:49 | 2.3694 | 2.398 | 2016-09-19 08:50 | 2.3509 | 2.3786 | 2016-09-19 08:49 | 2.3788 | 2.4077 | 2016-09-19 08:50 | 2.3601 | 2.3883 | 0.42039 | 0.12128 | 0.42372 | 0.38387
76 | 43.3 | 143.2 | 19 | 8 | 52 | 0 | 9 | 16 | H | 83 | IZO | 71.6231 | 42.3418 | 131.2899 | 273.0083 | 289.5233 | 273.1442 | 260.2683 | 268.9885 | 0.432 | 0.22564 | 0.43531 | 0.10441 | 0.4386 | 0.48827 | 2016-09-19 08:52 | 2.3148 | 2.3411 | 2016-09-19 08:53 | 2.2972 | 2.3228 | 2016-09-19 08:54 | 2.28 | 2.3048 | 2016-09-19 08:52 | 2.306 | 2.3319 | 2016-09-19 08:54 | 2.2886 | 2.3138 | 0.43366 | 0.12124 | 0.43696 | 0.38386
73.5 | 41.9 | 139.1 | 19 | 8 | 56 | 0 | 9 | 16 | H | 83 | IZO | 69.3191 | 40.9799 | 127.3791 | 272.1022 | 288.2086 | 272.937 | 259.6747 | 269.0623 | 0.44518 | 0.22489 | 0.44845 | 0.10393 | 0.45172 | 0.4879 | 2016-09-19 08:56 | 2.2463 | 2.2698 | 2016-09-19 08:57 | 2.2299 | 2.2528 | 2016-09-19 08:58 | 2.2137 | 2.236 | 2016-09-19 08:56 | 2.2381 | 2.2613 | 2016-09-19 08:58 | 2.2218 | 2.2444 | 0.44682 | 0.12096 | 0.45009 | 0.38397
71.3 | 40.8 | 135.3 | 19 | 9 | 0 | 0 | 9 | 16 | H | 83 | IZO | 67.292 | 39.9048 | 123.7563 | 271.8972 | 289.43 | 272.8207 | 258.3623 | 268.6091 | 0.45824 | 0.22472 | 0.46149 | 0.10437 | 0.46473 | 0.48769 | 2016-09-19 09:00 | 2.1823 | 2.2034 | 2016-09-19 09:01 | 2.1669 | 2.1875 | 2016-09-19 09:02 | 2.1518 | 2.1719 | 2016-09-19 09:00 | 2.1746 | 2.1954 | 2016-09-19 09:02 | 2.1593 | 2.1797 | 0.45986 | 0.12035 | 0.46311 | 0.38332
69.4 | 39.6 | 131.8 | 19 | 9 | 4 | 0 | 9 | 16 | H | 83 | IZO | 65.543 | 38.7269 | 120.4106 | 272.5359 | 288.3759 | 272.8332 | 260.3085 | 268.8921 | 0.47118 | 0.22525 | 0.4744 | 0.10399 | 0.47761 | 0.48771 | 2016-09-19 09:04 | 2.1223 | 2.1414 | 2016-09-19 09:05 | 2.1079 | 2.1265 | 2016-09-19 09:06 | 2.0938 | 2.1119 | 2016-09-19 09:05 | 2.1151 | 2.1339 | 2016-09-19 09:06 | 2.1008 | 2.1191 | 0.47279 | 0.12126 | 0.47601 | 0.38372
67.6 | 38.6 | 128.7 | 19 | 9 | 8 | 0 | 9 | 16 | H | 83 | IZO | 63.8857 | 37.7407 | 117.4517 | 273.0622 | 288.9261 | 273.3391 | 260.8137 | 269.3876 | 0.484 | 0.22569 | 0.48719 | 0.10419 | 0.49037 | 0.48862 | 2016-09-19 09:08 | 2.0661 | 2.0833 | 2016-09-19 09:09 | 2.0526 | 2.0694 | 2016-09-19 09:10 | 2.0393 | 2.0557 | 2016-09-19 09:08 | 2.0593 | 2.0763 | 2016-09-19 09:10 | 2.0459 | 2.0625 | 0.4856 | 0.1215 | 0.48878 | 0.38443
65.8 | 37.6 | 125.4 | 19 | 9 | 12 | 0 | 9 | 16 | H | 83 | IZO | 62.2258 | 36.7508 | 114.3079 | 272.9563 | 288.579 | 272.8158 | 260.8923 | 268.8205 | 0.49669 | 0.2256 | 0.49985 | 0.10406 | 0.50299 | 0.48768 | 2016-09-19 09:12 | 2.0133 | 2.0289 | 2016-09-19 09:13 | 2.0006 | 2.0158 | 2016-09-19 09:14 | 1.9881 | 2.0029 | 2016-09-19 09:12 | 2.0069 | 2.0223 | 2016-09-19 09:14 | 1.9943 | 2.0093 | 0.49827 | 0.12153 | 0.50142 | 0.38361
64 | 36.9 | 122.7 | 19 | 9 | 16 | 0 | 9 | 16 | H | 83 | IZO | 60.5621 | 36.0557 | 111.7494 | 272.2144 | 291.5825 | 273.4489 | 257.2493 | 268.8552 | 0.50925 | 0.22499 | 0.51237 | 0.10515 | 0.51548 | 0.48881 | 2016-09-19 09:16 | 1.9637 | 1.9778 | 2016-09-19 09:17 | 1.9517 | 1.9655 | 2016-09-19 09:18 | 1.9399 | 1.9534 | 2016-09-19 09:16 | 1.9577 | 1.9716 | 2016-09-19 09:18 | 1.9458 | 1.9594 | 0.51081 | 0.11984 | 0.51392 | 0.38366
62.7 | 36 | 119.8 | 19 | 9 | 20 | 0 | 9 | 16 | H | 83 | IZO | 59.3593 | 35.1597 | 109.0228 | 273.772 | 291.0335 | 273.1401 | 260.4356 | 268.6078 | 0.52167 | 0.22627 | 0.52475 | 0.10495 | 0.52783 | 0.48826 | 2016-09-19 09:20 | 1.9169 | 1.9297 | 2016-09-19 09:21 | 1.9057 | 1.9182 | 2016-09-19 09:22 | 1.8946 | 1.9068 | 2016-09-19 09:20 | 1.9113 | 1.9239 | 2016-09-19 09:22 | 1.9001 | 1.9125 | 0.52321 | 0.12132 | 0.52629 | 0.38331
61.2 | 35.3 | 117 | 19 | 9 | 24 | 0 | 9 | 16 | H | 83 | IZO | 57.9693 | 34.4613 | 106.3964 | 273.6546 | 292.6639 | 272.6686 | 258.9637 | 267.6056 | 0.53394 | 0.22618 | 0.53699 | 0.10554 | 0.54003 | 0.48742 | 2016-09-19 09:24 | 1.8729 | 1.8845 | 2016-09-19 09:25 | 1.8622 | 1.8736 | 2016-09-19 09:26 | 1.8518 | 1.8629 | 2016-09-19 09:24 | 1.8675 | 1.879 | 2016-09-19 09:26 | 1.857 | 1.8682 | 0.53547 | 0.12064 | 0.53851 | 0.38188
59.6 | 34.4 | 114.7 | 19 | 9 | 28 | 0 | 9 | 16 | H | 83 | IZO | 56.4835 | 33.5614 | 104.2343 | 272.3854 | 290.6062 | 273.1841 | 258.3034 | 268.7718 | 0.54607 | 0.22513 | 0.54908 | 0.1048 | 0.55208 | 0.48834 | 2016-09-19 09:28 | 1.8313 | 1.8419 | 2016-09-19 09:29 | 1.8212 | 1.8316 | 2016-09-19 09:30 | 1.8113 | 1.8215 | 2016-09-19 09:28 | 1.8262 | 1.8367 | 2016-09-19 09:30 | 1.8163 | 1.8265 | 0.54758 | 0.12033 | 0.55058 | 0.38354
58.4 | 33.8 | 112.1 | 19 | 9 | 32 | 0 | 9 | 16 | H | 83 | IZO | 55.3665 | 32.9605 | 101.7915 | 273.0662 | 292.4495 | 272.4323 | 258.083 | 267.3647 | 0.55805 | 0.22569 | 0.56102 | 0.10546 | 0.56398 | 0.48699 | 2016-09-19 09:32 | 1.792 | 1.8017 | 2016-09-19 09:33 | 1.7825 | 1.792 | 2016-09-19 09:34 | 1.7731 | 1.7824 | 2016-09-19 09:32 | 1.7872 | 1.7968 | 2016-09-19 09:34 | 1.7778 | 1.7872 | 0.55953 | 0.12023 | 0.5625 | 0.38153
57.2 | 33.2 | 110.1 | 19 | 9 | 36 | 0 | 9 | 16 | H | 83 | IZO | 54.2483 | 32.3588 | 99.9188 | 273.3097 | 293.7352 | 273.1017 | 257.5185 | 267.8787 | 0.56986 | 0.22589 | 0.57279 | 0.10592 | 0.57571 | 0.48819 | 2016-09-19 09:36 | 1.7548 | 1.7637 | 2016-09-19 09:37 | 1.7458 | 1.7545 | 2016-09-19 09:38 | 1.737 | 1.7454 | 2016-09-19 09:36 | 1.7503 | 1.7591 | 2016-09-19 09:38 | 1.7414 | 1.75 | 0.57133 | 0.11997 | 0.57425 | 0.38227
56.2 | 32.6 | 107.9 | 19 | 9 | 40 | 0 | 9 | 16 | H | 83 | IZO | 53.3153 | 31.7563 | 97.8697 | 274.4298 | 294.4663 | 272.8553 | 258.9387 | 267.3856 | 0.58152 | 0.22682 | 0.58441 | 0.10619 | 0.58729 | 0.48775 | 2016-09-19 09:40 | 1.7196 | 1.7277 | 2016-09-19 09:41 | 1.7111 | 1.7191 | 2016-09-19 09:42 | 1.7027 | 1.7105 | 2016-09-19 09:40 | 1.7154 | 1.7234 | 2016-09-19 09:42 | 1.7069 | 1.7148 | 0.58297 | 0.12063 | 0.58585 | 0.38156
54.8 | 32 | 106.1 | 19 | 9 | 44 | 0 | 9 | 16 | H | 83 | IZO | 52.0062 | 31.1529 | 96.1962 | 272.4808 | 294.6457 | 273.5104 | 255.3412 | 268.1612 | 0.59301 | 0.22521 | 0.59586 | 0.10625 | 0.5987 | 0.48892 | 2016-09-19 09:44 | 1.6863 | 1.6937 | 2016-09-19 09:45 | 1.6782 | 1.6855 | 2016-09-19 09:46 | 1.6703 | 1.6774 | 2016-09-19 09:44 | 1.6823 | 1.6896 | 2016-09-19 09:46 | 1.6743 | 1.6814 | 0.59444 | 0.11895 | 0.59728 | 0.38267
54.1 | 31.5 | 104.2 | 19 | 9 | 48 | 0 | 9 | 16 | H | 83 | IZO | 51.3499 | 30.6493 | 94.4251 | 274.8358 | 295.9741 | 273.5275 | 258.49 | 267.8473 | 0.60434 | 0.22715 | 0.60714 | 0.10673 | 0.60993 | 0.48895 | 2016-09-19 09:48 | 1.6547 | 1.6615 | 2016-09-19 09:49 | 1.6471 | 1.6537 | 2016-09-19 09:50 | 1.6395 | 1.646 | 2016-09-19 09:48 | 1.6509 | 1.6576 | 2016-09-19 09:50 | 1.6433 | 1.6499 | 0.60574 | 0.12042 | 0.60854 | 0.38222
47.5 | 27.6 | 92.1 | 19 | 10 | 20 | 0 | 9 | 16 | H | 83 | IZO | 45.0998 | 26.6942 | 83.0492 | 275.1649 | 291.6863 | 273.4862 | 262.3878 | 268.8807 | 0.68837 | 0.22742 | 0.6908 | 0.10519 | 0.69322 | 0.48888 | 2016-09-19 10:20 | 1.4527 | 1.4562 | 2016-09-19 10:21 | 1.4476 | 1.451 | 2016-09-19 10:22 | 1.4426 | 1.4459 | 2016-09-19 10:20 | 1.4501 | 1.4536 | 2016-09-19 10:22 | 1.4451 | 1.4484 | 0.68959 | 0.12224 | 0.69201 | 0.38369
46.7 | 27.2 | 90.9 | 19 | 10 | 24 | 0 | 9 | 16 | H | 83 | IZO | 44.3328 | 26.2855 | 81.9257 | 273.96 | 290.876 | 273.5055 | 260.8771 | 269.1098 | 0.69801 | 0.22643 | 0.70039 | 0.10489 | 0.70275 | 0.48891 | 2016-09-19 10:24 | 1.4326 | 1.4359 | 2016-09-19 10:25 | 1.4278 | 1.4309 | 2016-09-19 10:26 | 1.423 | 1.4261 | 2016-09-19 10:24 | 1.4302 | 1.4334 | 2016-09-19 10:25 | 1.4254 | 1.4285 | 0.6992 | 0.12154 | 0.70157 | 0.38402
46.2 | 26.8 | 89.7 | 19 | 10 | 28 | 0 | 9 | 16 | H | 83 | IZO | 43.8519 | 25.8762 | 80.8061 | 274.9167 | 289.6879 | 273.327 | 263.4935 | 269.1866 | 0.70744 | 0.22722 | 0.70976 | 0.10446 | 0.71208 | 0.48859 | 2016-09-19 10:28 | 1.4136 | 1.4165 | 2016-09-19 10:29 | 1.4089 | 1.4118 | 2016-09-19 10:30 | 1.4043 | 1.4072 | 2016-09-19 10:28 | 1.4112 | 1.4142 | 2016-09-19 10:30 | 1.4066 | 1.4095 | 0.7086 | 0.12275 | 0.71092 | 0.38413
45.5 | 26.6 | 88.7 | 19 | 10 | 32 | 0 | 9 | 16 | H | 83 | IZO | 43.177 | 25.6713 | 79.8753 | 273.9723 | 292.2139 | 273.7 | 259.8625 | 269.0157 | 0.71666 | 0.22644 | 0.71893 | 0.10538 | 0.72119 | 0.48926 | 2016-09-19 10:32 | 1.3954 | 1.3981 | 2016-09-19 10:33 | 1.3909 | 1.3936 | 2016-09-19 10:34 | 1.3866 | 1.3892 | 2016-09-19 10:32 | 1.3931 | 1.3958 | 2016-09-19 10:34 | 1.3888 | 1.3914 | 0.7178 | 0.12106 | 0.72007 | 0.38388
45.1 | 26.4 | 87.9 | 19 | 10 | 36 | 0 | 9 | 16 | H | 83 | IZO | 42.7907 | 25.4663 | 79.132 | 275.3011 | 294.4818 | 274.6493 | 260.4641 | 269.6319 | 0.72567 | 0.22754 | 0.72789 | 0.10619 | 0.7301 | 0.49096 | 2016-09-19 10:36 | 1.378 | 1.3805 | 2016-09-19 10:37 | 1.3738 | 1.3763 | 2016-09-19 10:38 | 1.3697 | 1.3721 | 2016-09-19 10:36 | 1.3759 | 1.3784 | 2016-09-19 10:38 | 1.3717 | 1.3742 | 0.72678 | 0.12134 | 0.729 | 0.38476
44.6 | 26.1 | 86.9 | 19 | 10 | 40 | 0 | 9 | 16 | H | 83 | IZO | 42.307 | 25.1586 | 78.2039 | 275.5702 | 294.3971 | 274.6668 | 261.0068 | 269.6752 | 0.73447 | 0.22776 | 0.73664 | 0.10616 | 0.73879 | 0.49099 | 2016-09-19 10:40 | 1.3615 | 1.3638 | 2016-09-19 10:41 | 1.3575 | 1.3598 | 2016-09-19 10:42 | 1.3536 | 1.3558 | 2016-09-19 10:40 | 1.3595 | 1.3618 | 2016-09-19 10:41 | 1.3555 | 1.3578 | 0.73556 | 0.1216 | 0.73772 | 0.38483
43.9 | 25.8 | 86.1 | 19 | 10 | 44 | 0 | 9 | 16 | H | 83 | IZO | 41.6283 | 24.8505 | 77.4622 | 273.8731 | 294.0056 | 275.2785 | 258.2987 | 270.5405 | 0.74305 | 0.22636 | 0.74516 | 0.10602 | 0.74726 | 0.49208 | 2016-09-19 10:44 | 1.3458 | 1.3479 | 2016-09-19 10:45 | 1.342 | 1.3441 | 2016-09-19 10:46 | 1.3382 | 1.3403 | 2016-09-19 10:44 | 1.3439 | 1.346 | 2016-09-19 10:46 | 1.3401 | 1.3422 | 0.74411 | 0.12034 | 0.74622 | 0.38606
43.8 | 26.1 | 85.7 | 19 | 10 | 48 | 0 | 9 | 16 | H | 83 | IZO | 41.5311 | 25.1586 | 77.0915 | 277.2134 | 306.1911 | 277.2931 | 254.7923 | 269.9849 | 0.75142 | 0.22912 | 0.75347 | 0.11042 | 0.75552 | 0.49568 | 2016-09-19 10:48 | 1.3308 | 1.3328 | 2016-09-19 10:49 | 1.3272 | 1.3291 | 2016-09-19 10:50 | 1.3236 | 1.3255 | 2016-09-19 10:48 | 1.329 | 1.3309 | 2016-09-19 10:50 | 1.3254 | 1.3273 | 0.75245 | 0.1187 | 0.7545 | 0.38527
43.5 | 25.9 | 84.8 | 19 | 10 | 52 | 0 | 9 | 16 | H | 83 | IZO | 41.2396 | 24.9532 | 76.2578 | 278.6352 | 307.5177 | 277.1994 | 256.2875 | 269.5324 | 0.75956 | 0.23029 | 0.76156 | 0.11089 | 0.76354 | 0.49552 | 2016-09-19 10:52 | 1.3166 | 1.3184 | 2016-09-19 10:53 | 1.3131 | 1.3149 | 2016-09-19 10:54 | 1.3097 | 1.3114 | 2016-09-19 10:52 | 1.3148 | 1.3166 | 2016-09-19 10:54 | 1.3114 | 1.3131 | 0.76056 | 0.1194 | 0.76255 | 0.38462
42.8 | 25.6 | 84.1 | 19 | 10 | 56 | 0 | 9 | 16 | H | 83 | IZO | 40.5579 | 24.6449 | 75.61 | 276.2652 | 306.4004 | 277.7378 | 252.9477 | 270.4893 | 0.76747 | 0.22833 | 0.76942 | 0.11049 | 0.77135 | 0.49648 | 2016-09-19 10:56 | 1.303 | 1.3046 | 2016-09-19 10:57 | 1.2997 | 1.3013 | 2016-09-19 10:58 | 1.2964 | 1.298 | 2016-09-19 10:56 | 1.3013 | 1.303 | 2016-09-19 10:57 | 1.2981 | 1.2997 | 0.76845 | 0.11784 | 0.77039 | 0.38599
42.5 | 25.4 | 83.2 | 19 | 11 | 0 | 0 | 9 | 16 | H | 83 | IZO | 40.2652 | 24.4393 | 74.7787 | 277.3022 | 307.2046 | 277.3241 | 254.1646 | 269.7679 | 0.77516 | 0.22919 | 0.77705 | 0.11078 | 0.77892 | 0.49574 | 2016-09-19 11:00 | 1.29 | 1.2916 | 2016-09-19 11:01 | 1.2869 | 1.2884 | 2016-09-19 11:02 | 1.2838 | 1.2853 | 2016-09-19 11:01 | 1.2885 | 1.29 | 2016-09-19 11:02 | 1.2854 | 1.2868 | 0.77611 | 0.11841 | 0.77799 | 0.38496
41.4 | 24.9 | 82.4 | 19 | 11 | 4 | 0 | 9 | 16 | H | 83 | IZO | 39.1884 | 23.9245 | 74.0408 | 270.7493 | 301.0423 | 277.155 | 247.3089 | 271.1135 | 0.78263 | 0.22377 | 0.78445 | 0.10856 | 0.78627 | 0.49544 | 2016-09-19 11:04 | 1.2778 | 1.2792 | 2016-09-19 11:05 | 1.2748 | 1.2761 | 2016-09-19 11:06 | 1.2718 | 1.2732 | 2016-09-19 11:04 | 1.2763 | 1.2776 | 2016-09-19 11:06 | 1.2733 | 1.2747 | 0.78354 | 0.11522 | 0.78536 | 0.38688
41.7 | 24.8 | 81.5 | 19 | 11 | 8 | 0 | 9 | 16 | H | 83 | IZO | 39.4826 | 23.8215 | 73.2113 | 276.9964 | 303.5549 | 276.4228 | 256.4467 | 269.5614 | 0.78985 | 0.22894 | 0.79163 | 0.10947 | 0.79338 | 0.49413 | 2016-09-19 11:08 | 1.2661 | 1.2673 | 2016-09-19 11:09 | 1.2632 | 1.2645 | 2016-09-19 11:10 | 1.2604 | 1.2617 | 2016-09-19 11:08 | 1.2646 | 1.2659 | 2016-09-19 11:10 | 1.2618 | 1.2631 | 0.79074 | 0.11947 | 0.79251 | 0.38466
41.3 | 24.4 | 80.8 | 19 | 11 | 12 | 0 | 9 | 16 | H | 83 | IZO | 39.0902 | 23.409 | 72.5661 | 276.5627 | 299.0182 | 276.3547 | 259.1886 | 270.6226 | 0.79685 | 0.22858 | 0.79856 | 0.10783 | 0.80026 | 0.49401 | 2016-09-19 11:12 | 1.2549 | 1.2561 | 2016-09-19 11:13 | 1.2523 | 1.2534 | 2016-09-19 11:14 | 1.2496 | 1.2507 | 2016-09-19 11:12 | 1.2536 | 1.2548 | 2016-09-19 11:13 | 1.2509 | 1.2521 | 0.79771 | 0.12075 | 0.79941 | 0.38618
40.3 | 24.1 | 80.3 | 19 | 11 | 16 | 0 | 9 | 16 | H | 83 | IZO | 38.1061 | 23.0992 | 72.1051 | 270.1975 | 296.4663 | 276.9726 | 249.8714 | 272.0415 | 0.80361 | 0.22332 | 0.80526 | 0.10691 | 0.8069 | 0.49511 | 2016-09-19 11:16 | 1.2444 | 1.2455 | 2016-09-19 11:17 | 1.2418 | 1.2429 | 2016-09-19 11:18 | 1.2393 | 1.2403 | 2016-09-19 11:17 | 1.2431 | 1.2442 | 2016-09-19 11:18 | 1.2406 | 1.2416 | 0.80444 | 0.11641 | 0.80608 | 0.3882
40.5 | 24 | 79.7 | 19 | 11 | 20 | 0 | 9 | 16 | H | 83 | IZO | 38.3033 | 22.9959 | 71.5516 | 275.1433 | 298.2937 | 277.0391 | 257.231 | 271.663 | 0.81013 | 0.22741 | 0.81172 | 0.10757 | 0.8133 | 0.49523 | 2016-09-19 11:20 | 1.2344 | 1.2354 | 2016-09-19 11:21 | 1.232 | 1.2329 | 2016-09-19 11:22 | 1.2296 | 1.2305 | 2016-09-19 11:20 | 1.2332 | 1.2341 | 2016-09-19 11:22 | 1.2308 | 1.2317 | 0.81093 | 0.11984 | 0.81251 | 0.38766
40.2 | 23.9 | 79 | 19 | 11 | 24 | 0 | 9 | 16 | H | 83 | IZO | 38.0074 | 22.8926 | 70.9052 | 275.1371 | 299.9294 | 276.5438 | 255.9536 | 270.6293 | 0.81641 | 0.2274 | 0.81794 | 0.10816 | 0.81946 | 0.49434 | 2016-09-19 11:24 | 1.2249 | 1.2258 | 2016-09-19 11:25 | 1.2226 | 1.2235 | 2016-09-19 11:26 | 1.2203 | 1.2212 | 2016-09-19 11:24 | 1.2237 | 1.2246 | 2016-09-19 11:26 | 1.2214 | 1.2223 | 0.81718 | 0.11924 | 0.8187 | 0.38619

## IdD083_star
This Page uses the lamp tests and the langley plot results in the limited mu range from sheet IdD083 | [21 empty] | the correction values from ID083 are applied to the NC
Dobson 083 at IZO 2016.Sep.19 Langley Plot Results | [15 empty] | 109
SLC083Na | -0.9 | SLC083Nc | -0.7 | [9 empty] | SLC083Nd | -0.7
D083 LP adj. A:0.55 C:-0.15 D:0.12 Nad:0.43 | [20 empty] | LPC_C_83 |  | LPC_D_83 |  | LPC_A_83
[15 empty] | 268.9581 | 272.5002 | 267.3461 | 266.2646 | 266.0256 | All Obs | -0.15348 |  | 0.12348 |  | 0.54874 | 0.42526 | -0.27696
D083RC | D083RD | D083RA | day | h | m | s | m | year |  |  |  | D083NC | D083ND | D083NA | D083_TO3_C* | D083_TO3_D* | D083_TO3_A* | D083_TO3_CD* | D083_TO3_AD* | 1/mu.C | P*C_83 | 1/Mu.D | P*D_83 | 1/Mu.A | P*A_83 | lst.C | Mu.C | AirM.C | lst.D | Mu.D | AirM.D | lst.A | Mu.A | AirM.A | lst.CD | Mu.CD | AirM.CD | lst.AD | Mu.AD | AirM.AD
138 | 79.4 | 245.8 | 19 | 7 | 56 | 0 | 9 | 16 | H | 83 | IZO | 128.5576 | 74.5847 | 230.8505 | 268.2594 | 274.2967 | 268.196 | 263.8567 | 266.5657 | 0.23991 | 0.22172 | 0.24338 | 0.098914 | 0.24685 | 0.47942 | 2016-09-19 07:56 | 4.1682 | 4.3631 | 2016-09-19 07:57 | 4.1089 | 4.2947 | 2016-09-19 07:58 | 4.0511 | 4.2284 | 2016-09-19 07:56 | 4.1383 | 4.3286 | 2016-09-19 07:57 | 4.0798 | 4.2613
130.3 | 74.9 | 233.7 | 19 | 8 | 0 | 0 | 9 | 16 | H | 83 | IZO | 121.3117 | 70.4611 | 218.3978 | 268.1766 | 274.8103 | 268.2378 | 263.2676 | 266.5023 | 0.25379 | 0.22165 | 0.25727 | 0.0991 | 0.26075 | 0.4795 | 2016-09-19 08:00 | 3.9402 | 4.1019 | 2016-09-19 08:01 | 3.887 | 4.0415 | 2016-09-19 08:02 | 3.8351 | 3.9829 | 2016-09-19 08:01 | 3.9134 | 4.0715 | 2016-09-19 08:02 | 3.8609 | 4.012
123.5 | 70.6 | 222.9 | 19 | 8 | 4 | 0 | 9 | 16 | H | 83 | IZO | 114.9698 | 66.5312 | 207.4282 | 268.5392 | 273.4602 | 268.6071 | 264.9206 | 267.3173 | 0.26771 | 0.22195 | 0.27119 | 0.098613 | 0.27467 | 0.48016 | 2016-09-19 08:04 | 3.7354 | 3.8708 | 2016-09-19 08:05 | 3.6875 | 3.8172 | 2016-09-19 08:06 | 3.6407 | 3.765 | 2016-09-19 08:04 | 3.7113 | 3.8438 | 2016-09-19 08:06 | 3.6639 | 3.7909
117.1 | 66.9 | 212.5 | 19 | 8 | 8 | 0 | 9 | 16 | H | 83 | IZO | 109.0215 | 63.1589 | 197.0367 | 268.0495 | 273.3452 | 268.1731 | 264.1125 | 266.8115 | 0.28164 | 0.22154 | 0.28512 | 0.098571 | 0.2886 | 0.47938 | 2016-09-19 08:08 | 3.5507 | 3.6651 | 2016-09-19 08:09 | 3.5073 | 3.6171 | 2016-09-19 08:10 | 3.465 | 3.5704 | 2016-09-19 08:08 | 3.5289 | 3.6409 | 2016-09-19 08:10 | 3.486 | 3.5936
111.4 | 63.6 | 203.2 | 19 | 8 | 12 | 0 | 9 | 16 | H | 83 | IZO | 103.7432 | 60.1366 | 187.8367 | 267.8861 | 273.3881 | 268.0977 | 263.7665 | 266.7135 | 0.29556 | 0.22141 | 0.29904 | 0.098587 | 0.30252 | 0.47925 | 2016-09-19 08:12 | 3.3834 | 3.4808 | 2016-09-19 08:13 | 3.344 | 3.4377 | 2016-09-19 08:14 | 3.3056 | 3.3957 | 2016-09-19 08:12 | 3.3636 | 3.4591 | 2016-09-19 08:13 | 3.3247 | 3.4165
106.6 | 60.7 | 194.9 | 19 | 8 | 16 | 0 | 9 | 16 | H | 83 | IZO | 99.295 | 57.4687 | 179.6664 | 268.9806 | 274.0368 | 268.3654 | 265.1874 | 266.8908 | 0.30947 | 0.22231 | 0.31294 | 0.098821 | 0.31641 | 0.47972 | 2016-09-19 08:16 | 3.2314 | 3.3148 | 2016-09-19 08:17 | 3.1955 | 3.2759 | 2016-09-19 08:18 | 3.1605 | 3.2379 | 2016-09-19 08:17 | 3.2133 | 3.2953 | 2016-09-19 08:18 | 3.1779 | 3.2568
102 | 58 | 187.1 | 19 | 8 | 20 | 0 | 9 | 16 | H | 83 | IZO | 95.023 | 54.9713 | 171.9817 | 269.191 | 274.0373 | 268.2066 | 265.545 | 266.6968 | 0.32334 | 0.22249 | 0.3268 | 0.098821 | 0.33026 | 0.47944 | 2016-09-19 08:20 | 3.0927 | 3.1647 | 2016-09-19 08:21 | 3.0599 | 3.1294 | 2016-09-19 08:22 | 3.0279 | 3.0949 | 2016-09-19 08:20 | 3.0762 | 3.147 | 2016-09-19 08:22 | 3.0438 | 3.1121
97.8 | 55.7 | 180.3 | 19 | 8 | 24 | 0 | 9 | 16 | H | 83 | IZO | 91.1182 | 52.8351 | 165.2719 | 269.3841 | 275.3925 | 268.6737 | 264.8257 | 266.9435 | 0.33717 | 0.22265 | 0.34062 | 0.09931 | 0.34407 | 0.48028 | 2016-09-19 08:24 | 2.9659 | 3.0283 | 2016-09-19 08:25 | 2.9358 | 2.9962 | 2016-09-19 08:26 | 2.9064 | 2.9647 | 2016-09-19 08:24 | 2.9508 | 3.0122 | 2016-09-19 08:26 | 2.921 | 2.9804
93.7 | 53.4 | 173.6 | 19 | 8 | 28 | 0 | 9 | 16 | H | 83 | IZO | 87.3179 | 50.6861 | 158.7197 | 268.6298 | 274.848 | 268.3557 | 263.8983 | 266.6862 | 0.35095 | 0.22202 | 0.35438 | 0.099113 | 0.35781 | 0.47971 | 2016-09-19 08:28 | 2.8494 | 2.9039 | 2016-09-19 08:29 | 2.8218 | 2.8745 | 2016-09-19 08:30 | 2.7948 | 2.8458 | 2016-09-19 08:28 | 2.8355 | 2.8892 | 2016-09-19 08:29 | 2.8082 | 2.8601
90.4 | 51.3 | 167.6 | 19 | 8 | 32 | 0 | 9 | 16 | H | 83 | IZO | 84.2597 | 48.7084 | 152.8539 | 269.796 | 274.2749 | 268.3887 | 266.4016 | 266.8752 | 0.36466 | 0.22299 | 0.36808 | 0.098906 | 0.37149 | 0.47977 | 2016-09-19 08:32 | 2.7423 | 2.7901 | 2016-09-19 08:33 | 2.7168 | 2.7631 | 2016-09-19 08:34 | 2.6918 | 2.7367 | 2016-09-19 08:33 | 2.7295 | 2.7765 | 2016-09-19 08:34 | 2.7043 | 2.7498
86.9 | 49.3 | 161.7 | 19 | 8 | 36 | 0 | 9 | 16 | H | 83 | IZO | 81.0135 | 46.8126 | 147.0806 | 268.9965 | 272.9298 | 267.6522 | 266.0167 | 266.295 | 0.3783 | 0.22233 | 0.3817 | 0.098421 | 0.3851 | 0.47845 | 2016-09-19 08:36 | 2.6434 | 2.6855 | 2016-09-19 08:37 | 2.6198 | 2.6606 | 2016-09-19 08:38 | 2.5968 | 2.6363 | 2016-09-19 08:36 | 2.6316 | 2.673 | 2016-09-19 08:38 | 2.6082 | 2.6484
83.9 | 47.7 | 156.7 | 19 | 8 | 40 | 0 | 9 | 16 | H | 83 | IZO | 78.2413 | 45.287 | 142.2304 | 269.2805 | 274.0727 | 268.0218 | 265.6294 | 266.4714 | 0.39187 | 0.22256 | 0.39524 | 0.098834 | 0.39862 | 0.47911 | 2016-09-19 08:40 | 2.5519 | 2.5891 | 2016-09-19 08:41 | 2.5301 | 2.5662 | 2016-09-19 08:42 | 2.5087 | 2.5437 | 2016-09-19 08:40 | 2.5409 | 2.5776 | 2016-09-19 08:42 | 2.5193 | 2.5549
81.1 | 46.1 | 152 | 19 | 8 | 44 | 0 | 9 | 16 | H | 83 | IZO | 75.6649 | 43.7529 | 137.6755 | 269.5217 | 273.8349 | 268.2626 | 266.2355 | 266.8349 | 0.40534 | 0.22276 | 0.4087 | 0.098748 | 0.41204 | 0.47954 | 2016-09-19 08:44 | 2.4671 | 2.5001 | 2016-09-19 08:45 | 2.4468 | 2.4789 | 2016-09-19 08:46 | 2.4269 | 2.4581 | 2016-09-19 08:44 | 2.4569 | 2.4894 | 2016-09-19 08:46 | 2.4368 | 2.4684
78.4 | 44.6 | 147.4 | 19 | 8 | 48 | 0 | 9 | 16 | H | 83 | IZO | 73.181 | 42.3072 | 133.2314 | 269.2952 | 273.3495 | 268.0044 | 266.2046 | 266.6357 | 0.41872 | 0.22257 | 0.42205 | 0.098573 | 0.42538 | 0.47908 | 2016-09-19 08:48 | 2.3882 | 2.4176 | 2016-09-19 08:49 | 2.3694 | 2.398 | 2016-09-19 08:50 | 2.3509 | 2.3786 | 2016-09-19 08:49 | 2.3788 | 2.4077 | 2016-09-19 08:50 | 2.3601 | 2.3883
76 | 43.3 | 143.2 | 19 | 8 | 52 | 0 | 9 | 16 | H | 83 | IZO | 70.9698 | 41.0485 | 129.2155 | 269.5938 | 273.9116 | 268.0545 | 266.2949 | 266.5578 | 0.432 | 0.22282 | 0.43531 | 0.098775 | 0.4386 | 0.47917 | 2016-09-19 08:52 | 2.3148 | 2.3411 | 2016-09-19 08:53 | 2.2972 | 2.3228 | 2016-09-19 08:54 | 2.28 | 2.3048 | 2016-09-19 08:52 | 2.306 | 2.3319 | 2016-09-19 08:54 | 2.2886 | 2.3138
73.5 | 41.9 | 139.1 | 19 | 8 | 56 | 0 | 9 | 16 | H | 83 | IZO | 68.6659 | 39.6866 | 125.3047 | 268.5836 | 272.1254 | 267.695 | 265.8815 | 266.5601 | 0.44518 | 0.22199 | 0.44845 | 0.098131 | 0.45172 | 0.47853 | 2016-09-19 08:56 | 2.2463 | 2.2698 | 2016-09-19 08:57 | 2.2299 | 2.2528 | 2016-09-19 08:58 | 2.2137 | 2.236 | 2016-09-19 08:56 | 2.2381 | 2.2613 | 2016-09-19 08:58 | 2.2218 | 2.2444
71.3 | 40.8 | 135.3 | 19 | 9 | 0 | 0 | 9 | 16 | H | 83 | IZO | 66.6388 | 38.6116 | 121.6819 | 268.2753 | 272.8794 | 267.4278 | 264.7477 | 266.0359 | 0.45824 | 0.22173 | 0.46149 | 0.098403 | 0.46473 | 0.47805 | 2016-09-19 09:00 | 2.1823 | 2.2034 | 2016-09-19 09:01 | 2.1669 | 2.1875 | 2016-09-19 09:02 | 2.1518 | 2.1719 | 2016-09-19 09:00 | 2.1746 | 2.1954 | 2016-09-19 09:02 | 2.1593 | 2.1797
69.4 | 39.6 | 131.8 | 19 | 9 | 4 | 0 | 9 | 16 | H | 83 | IZO | 64.8897 | 37.4336 | 118.3362 | 268.8117 | 271.3621 | 267.2909 | 266.8708 | 266.2487 | 0.47118 | 0.22217 | 0.4744 | 0.097856 | 0.47761 | 0.4778 | 2016-09-19 09:04 | 2.1223 | 2.1414 | 2016-09-19 09:05 | 2.1079 | 2.1265 | 2016-09-19 09:06 | 2.0938 | 2.1119 | 2016-09-19 09:05 | 2.1151 | 2.1339 | 2016-09-19 09:06 | 2.1008 | 2.1191
67.6 | 38.6 | 128.7 | 19 | 9 | 8 | 0 | 9 | 16 | H | 83 | IZO | 63.2324 | 36.4475 | 115.3773 | 269.2367 | 271.4537 | 267.6487 | 267.5513 | 266.6747 | 0.484 | 0.22252 | 0.48719 | 0.097889 | 0.49037 | 0.47844 | 2016-09-19 09:08 | 2.0661 | 2.0833 | 2016-09-19 09:09 | 2.0526 | 2.0694 | 2016-09-19 09:10 | 2.0393 | 2.0557 | 2016-09-19 09:08 | 2.0593 | 2.0763 | 2016-09-19 09:10 | 2.0459 | 2.0625
65.8 | 37.6 | 125.4 | 19 | 9 | 12 | 0 | 9 | 16 | H | 83 | IZO | 61.5725 | 35.4575 | 112.2335 | 269.0305 | 270.6527 | 266.9789 | 267.8033 | 266.0387 | 0.49669 | 0.22235 | 0.49985 | 0.0976 | 0.50299 | 0.47725 | 2016-09-19 09:12 | 2.0133 | 2.0289 | 2016-09-19 09:13 | 2.0006 | 2.0158 | 2016-09-19 09:14 | 1.9881 | 2.0029 | 2016-09-19 09:12 | 2.0069 | 2.0223 | 2016-09-19 09:14 | 1.9943 | 2.0093
64 | 36.9 | 122.7 | 19 | 9 | 16 | 0 | 9 | 16 | H | 83 | IZO | 59.9089 | 34.7624 | 109.675 | 268.1894 | 273.207 | 267.4671 | 264.3319 | 266.0054 | 0.50925 | 0.22166 | 0.51237 | 0.098521 | 0.51548 | 0.47812 | 2016-09-19 09:16 | 1.9637 | 1.9778 | 2016-09-19 09:17 | 1.9517 | 1.9655 | 2016-09-19 09:18 | 1.9399 | 1.9534 | 2016-09-19 09:16 | 1.9577 | 1.9716 | 2016-09-19 09:18 | 1.9458 | 1.9594
62.7 | 36 | 119.8 | 19 | 9 | 20 | 0 | 9 | 16 | H | 83 | IZO | 58.706 | 33.8665 | 106.9484 | 269.6488 | 272.2139 | 267.015 | 267.6878 | 265.6907 | 0.52167 | 0.22287 | 0.52475 | 0.098163 | 0.52783 | 0.47731 | 2016-09-19 09:20 | 1.9169 | 1.9297 | 2016-09-19 09:21 | 1.9057 | 1.9182 | 2016-09-19 09:22 | 1.8946 | 1.9068 | 2016-09-19 09:20 | 1.9113 | 1.9239 | 2016-09-19 09:22 | 1.9001 | 1.9125
61 | 35.3 | 117 | 19 | 9 | 24 | 0 | 9 | 16 | H | 83 | IZO | 57.1305 | 33.168 | 104.322 | 268.2357 | 273.4054 | 266.4019 | 264.2571 | 264.6221 | 0.53394 | 0.2217 | 0.53699 | 0.098593 | 0.54003 | 0.47621 | 2016-09-19 09:24 | 1.8729 | 1.8845 | 2016-09-19 09:25 | 1.8622 | 1.8736 | 2016-09-19 09:26 | 1.8518 | 1.8629 | 2016-09-19 09:24 | 1.8675 | 1.879 | 2016-09-19 09:26 | 1.857 | 1.8682
59.6 | 34.4 | 114.7 | 19 | 9 | 28 | 0 | 9 | 16 | H | 83 | IZO | 55.8302 | 32.2681 | 102.16 | 268.0693 | 270.9142 | 266.7775 | 265.8888 | 265.7227 | 0.54607 | 0.22156 | 0.54908 | 0.097695 | 0.55208 | 0.47689 | 2016-09-19 09:28 | 1.8313 | 1.8419 | 2016-09-19 09:29 | 1.8212 | 1.8316 | 2016-09-19 09:30 | 1.8113 | 1.8215 | 2016-09-19 09:28 | 1.8262 | 1.8367 | 2016-09-19 09:30 | 1.8163 | 1.8265
58.4 | 33.8 | 112.1 | 19 | 9 | 32 | 0 | 9 | 16 | H | 83 | IZO | 54.7132 | 31.6672 | 99.7172 | 268.6555 | 272.3294 | 265.8877 | 265.8318 | 264.2508 | 0.55805 | 0.22204 | 0.56102 | 0.098205 | 0.56398 | 0.4753 | 2016-09-19 09:32 | 1.792 | 1.8017 | 2016-09-19 09:33 | 1.7825 | 1.792 | 2016-09-19 09:34 | 1.7731 | 1.7824 | 2016-09-19 09:32 | 1.7872 | 1.7968 | 2016-09-19 09:34 | 1.7778 | 1.7872
57.2 | 33.2 | 110.1 | 19 | 9 | 36 | 0 | 9 | 16 | H | 83 | IZO | 53.5951 | 31.0655 | 97.8444 | 268.8055 | 273.1928 | 266.4209 | 265.4286 | 264.701 | 0.56986 | 0.22217 | 0.57279 | 0.098516 | 0.57571 | 0.47625 | 2016-09-19 09:36 | 1.7548 | 1.7637 | 2016-09-19 09:37 | 1.7458 | 1.7545 | 2016-09-19 09:38 | 1.737 | 1.7454 | 2016-09-19 09:36 | 1.7503 | 1.7591 | 2016-09-19 09:38 | 1.7414 | 1.75
56.2 | 32.6 | 107.9 | 19 | 9 | 40 | 0 | 9 | 16 | H | 83 | IZO | 52.6621 | 30.463 | 95.7954 | 269.8335 | 273.5072 | 266.0402 | 267.0078 | 264.1449 | 0.58152 | 0.22302 | 0.58441 | 0.09863 | 0.58729 | 0.47557 | 2016-09-19 09:40 | 1.7196 | 1.7277 | 2016-09-19 09:41 | 1.7111 | 1.7191 | 2016-09-19 09:42 | 1.7027 | 1.7105 | 2016-09-19 09:40 | 1.7154 | 1.7234 | 2016-09-19 09:42 | 1.7069 | 1.7148
54.8 | 32 | 106.1 | 19 | 9 | 44 | 0 | 9 | 16 | H | 83 | IZO | 51.3529 | 29.8596 | 94.1218 | 267.7937 | 273.276 | 266.5629 | 263.567 | 264.8584 | 0.59301 | 0.22133 | 0.59586 | 0.098546 | 0.5987 | 0.4765 | 2016-09-19 09:44 | 1.6863 | 1.6937 | 2016-09-19 09:45 | 1.6782 | 1.6855 | 2016-09-19 09:46 | 1.6703 | 1.6774 | 2016-09-19 09:44 | 1.6823 | 1.6896 | 2016-09-19 09:46 | 1.6743 | 1.6814
54.1 | 31.5 | 104.2 | 19 | 9 | 48 | 0 | 9 | 16 | H | 83 | IZO | 50.6967 | 29.356 | 92.3507 | 270.0592 | 274.1999 | 266.4496 | 266.8701 | 264.4834 | 0.60434 | 0.2232 | 0.60714 | 0.098879 | 0.60993 | 0.4763 | 2016-09-19 09:48 | 1.6547 | 1.6615 | 2016-09-19 09:49 | 1.6471 | 1.6537 | 2016-09-19 09:50 | 1.6395 | 1.646 | 2016-09-19 09:48 | 1.6509 | 1.6576 | 2016-09-19 09:50 | 1.6433 | 1.6499
47.5 | 27.6 | 92.1 | 19 | 10 | 20 | 0 | 9 | 16 | H | 83 | IZO | 44.4465 | 25.4009 | 80.9748 | 269.7241 | 266.9116 | 265.4418 | 271.9119 | 265.0643 | 0.68837 | 0.22293 | 0.6908 | 0.096251 | 0.69322 | 0.4745 | 2016-09-19 10:20 | 1.4527 | 1.4562 | 2016-09-19 10:21 | 1.4476 | 1.451 | 2016-09-19 10:22 | 1.4426 | 1.4459 | 2016-09-19 10:20 | 1.4501 | 1.4536 | 2016-09-19 10:22 | 1.4451 | 1.4484
46.7 | 27.2 | 90.9 | 19 | 10 | 24 | 0 | 9 | 16 | H | 83 | IZO | 43.6795 | 24.9922 | 79.8513 | 268.4431 | 265.7576 | 265.3505 | 270.5322 | 265.2417 | 0.69801 | 0.22187 | 0.70039 | 0.095835 | 0.70275 | 0.47433 | 2016-09-19 10:24 | 1.4326 | 1.4359 | 2016-09-19 10:25 | 1.4278 | 1.4309 | 2016-09-19 10:26 | 1.423 | 1.4261 | 2016-09-19 10:24 | 1.4302 | 1.4334 | 2016-09-19 10:25 | 1.4254 | 1.4285
46.2 | 26.8 | 89.7 | 19 | 10 | 28 | 0 | 9 | 16 | H | 83 | IZO | 43.1986 | 24.5829 | 78.7317 | 269.3252 | 264.2332 | 265.0638 | 273.2767 | 265.2679 | 0.70744 | 0.2226 | 0.70976 | 0.095285 | 0.71208 | 0.47382 | 2016-09-19 10:28 | 1.4136 | 1.4165 | 2016-09-19 10:29 | 1.4089 | 1.4118 | 2016-09-19 10:30 | 1.4043 | 1.4072 | 2016-09-19 10:28 | 1.4112 | 1.4142 | 2016-09-19 10:30 | 1.4066 | 1.4095
45.5 | 26.6 | 88.7 | 19 | 10 | 32 | 0 | 9 | 16 | H | 83 | IZO | 42.5238 | 24.378 | 77.801 | 268.3079 | 266.4303 | 265.331 | 269.771 | 265.0476 | 0.71666 | 0.22176 | 0.71893 | 0.096078 | 0.72119 | 0.4743 | 2016-09-19 10:32 | 1.3954 | 1.3981 | 2016-09-19 10:33 | 1.3909 | 1.3936 | 2016-09-19 10:34 | 1.3866 | 1.3892 | 2016-09-19 10:32 | 1.3931 | 1.3958 | 2016-09-19 10:34 | 1.3888 | 1.3914
45.1 | 26.4 | 87.9 | 19 | 10 | 36 | 0 | 9 | 16 | H | 83 | IZO | 42.1374 | 24.173 | 77.0576 | 269.5655 | 268.3769 | 266.177 | 270.4949 | 265.6155 | 0.72567 | 0.2228 | 0.72789 | 0.09678 | 0.7301 | 0.47581 | 2016-09-19 10:36 | 1.378 | 1.3805 | 2016-09-19 10:37 | 1.3738 | 1.3763 | 2016-09-19 10:38 | 1.3697 | 1.3721 | 2016-09-19 10:36 | 1.3759 | 1.3784 | 2016-09-19 10:38 | 1.3717 | 1.3742
44.6 | 26.1 | 86.9 | 19 | 10 | 40 | 0 | 9 | 16 | H | 83 | IZO | 41.6537 | 23.8653 | 76.1295 | 269.7651 | 267.9786 | 266.0936 | 271.157 | 265.6118 | 0.73447 | 0.22296 | 0.73664 | 0.096636 | 0.73879 | 0.47566 | 2016-09-19 10:40 | 1.3615 | 1.3638 | 2016-09-19 10:41 | 1.3575 | 1.3598 | 2016-09-19 10:42 | 1.3536 | 1.3558 | 2016-09-19 10:40 | 1.3595 | 1.3618 | 2016-09-19 10:41 | 1.3555 | 1.3578
43.9 | 25.8 | 86.1 | 19 | 10 | 44 | 0 | 9 | 16 | H | 83 | IZO | 40.975 | 23.5572 | 75.3878 | 268.0001 | 267.2813 | 266.6069 | 268.5653 | 266.4312 | 0.74305 | 0.2215 | 0.74516 | 0.096385 | 0.74726 | 0.47658 | 2016-09-19 10:44 | 1.3458 | 1.3479 | 2016-09-19 10:45 | 1.342 | 1.3441 | 2016-09-19 10:46 | 1.3382 | 1.3403 | 2016-09-19 10:44 | 1.3439 | 1.346 | 2016-09-19 10:46 | 1.3401 | 1.3422
43.8 | 26.1 | 85.7 | 19 | 10 | 48 | 0 | 9 | 16 | H | 83 | IZO | 40.8779 | 23.8653 | 75.0172 | 271.2743 | 279.1688 | 268.5258 | 265.1722 | 265.8309 | 0.75142 | 0.22421 | 0.75347 | 0.10067 | 0.75552 | 0.48001 | 2016-09-19 10:48 | 1.3308 | 1.3328 | 2016-09-19 10:49 | 1.3272 | 1.3291 | 2016-09-19 10:50 | 1.3236 | 1.3255 | 2016-09-19 10:48 | 1.329 | 1.3309 | 2016-09-19 10:50 | 1.3254 | 1.3273
43.5 | 25.9 | 84.8 | 19 | 10 | 52 | 0 | 9 | 16 | H | 83 | IZO | 40.5863 | 23.6599 | 74.1834 | 272.6318 | 280.2055 | 268.3389 | 266.7778 | 265.335 | 0.75956 | 0.22533 | 0.76156 | 0.10105 | 0.76354 | 0.47968 | 2016-09-19 10:52 | 1.3166 | 1.3184 | 2016-09-19 10:53 | 1.3131 | 1.3149 | 2016-09-19 10:54 | 1.3097 | 1.3114 | 2016-09-19 10:52 | 1.3148 | 1.3166 | 2016-09-19 10:54 | 1.3114 | 1.3131
42.8 | 25.6 | 84.1 | 19 | 10 | 56 | 0 | 9 | 16 | H | 83 | IZO | 39.9047 | 23.3517 | 73.5356 | 270.1992 | 278.8063 | 268.7868 | 263.5451 | 266.2497 | 0.76747 | 0.22332 | 0.76942 | 0.10054 | 0.77135 | 0.48048 | 2016-09-19 10:56 | 1.303 | 1.3046 | 2016-09-19 10:57 | 1.2997 | 1.3013 | 2016-09-19 10:58 | 1.2964 | 1.298 | 2016-09-19 10:56 | 1.3013 | 1.303 | 2016-09-19 10:57 | 1.2981 | 1.2997
42.5 | 25.4 | 83.2 | 19 | 11 | 0 | 0 | 9 | 16 | H | 83 | IZO | 39.6119 | 23.146 | 72.7043 | 271.1754 | 279.3367 | 268.2851 | 264.8661 | 265.4873 | 0.77516 | 0.22413 | 0.77705 | 0.10073 | 0.77892 | 0.47958 | 2016-09-19 11:00 | 1.29 | 1.2916 | 2016-09-19 11:01 | 1.2869 | 1.2884 | 2016-09-19 11:02 | 1.2838 | 1.2853 | 2016-09-19 11:01 | 1.2885 | 1.29 | 2016-09-19 11:02 | 1.2854 | 1.2868
41.4 | 24.9 | 82.4 | 19 | 11 | 4 | 0 | 9 | 16 | H | 83 | IZO | 38.5351 | 22.6312 | 71.9664 | 264.5635 | 272.9089 | 268.0309 | 258.1113 | 266.7932 | 0.78263 | 0.21866 | 0.78445 | 0.098414 | 0.78627 | 0.47913 | 2016-09-19 11:04 | 1.2778 | 1.2792 | 2016-09-19 11:05 | 1.2748 | 1.2761 | 2016-09-19 11:06 | 1.2718 | 1.2732 | 2016-09-19 11:04 | 1.2763 | 1.2776 | 2016-09-19 11:06 | 1.2733 | 1.2747
41.7 | 24.8 | 81.5 | 19 | 11 | 8 | 0 | 9 | 16 | H | 83 | IZO | 38.8293 | 22.5282 | 71.1369 | 270.7535 | 275.1643 | 267.2161 | 267.3468 | 265.2027 | 0.78985 | 0.22378 | 0.79163 | 0.099227 | 0.79338 | 0.47767 | 2016-09-19 11:08 | 1.2661 | 1.2673 | 2016-09-19 11:09 | 1.2632 | 1.2645 | 2016-09-19 11:10 | 1.2604 | 1.2617 | 2016-09-19 11:08 | 1.2646 | 1.2659 | 2016-09-19 11:10 | 1.2618 | 1.2631
41.3 | 24.4 | 80.8 | 19 | 11 | 12 | 0 | 9 | 16 | H | 83 | IZO | 38.437 | 22.1157 | 70.4917 | 270.2645 | 270.3788 | 267.0682 | 270.1831 | 266.2268 | 0.79685 | 0.22337 | 0.79856 | 0.097502 | 0.80026 | 0.47741 | 2016-09-19 11:12 | 1.2549 | 1.2561 | 2016-09-19 11:13 | 1.2523 | 1.2534 | 2016-09-19 11:14 | 1.2496 | 1.2507 | 2016-09-19 11:12 | 1.2536 | 1.2548 | 2016-09-19 11:13 | 1.2509 | 1.2521
40.3 | 24.1 | 80.3 | 19 | 11 | 16 | 0 | 9 | 16 | H | 83 | IZO | 37.4528 | 21.806 | 70.0308 | 263.8459 | 267.5868 | 267.609 | 260.9572 | 267.6099 | 0.80361 | 0.21807 | 0.80526 | 0.096495 | 0.8069 | 0.47837 | 2016-09-19 11:16 | 1.2444 | 1.2455 | 2016-09-19 11:17 | 1.2418 | 1.2429 | 2016-09-19 11:18 | 1.2393 | 1.2403 | 2016-09-19 11:17 | 1.2431 | 1.2442 | 2016-09-19 11:18 | 1.2406 | 1.2416
40.5 | 24 | 79.7 | 19 | 11 | 20 | 0 | 9 | 16 | H | 83 | IZO | 37.65 | 21.7027 | 69.4773 | 268.7402 | 269.1824 | 267.6012 | 268.4046 | 267.197 | 0.81013 | 0.22211 | 0.81172 | 0.09707 | 0.8133 | 0.47836 | 2016-09-19 11:20 | 1.2344 | 1.2354 | 2016-09-19 11:21 | 1.232 | 1.2329 | 2016-09-19 11:22 | 1.2296 | 1.2305 | 2016-09-19 11:20 | 1.2332 | 1.2341 | 2016-09-19 11:22 | 1.2308 | 1.2317
40.2 | 23.9 | 79 | 19 | 11 | 24 | 0 | 9 | 16 | H | 83 | IZO | 37.3541 | 21.5993 | 68.8308 | 268.6843 | 270.5951 | 267.0345 | 267.2118 | 266.1301 | 0.81641 | 0.22207 | 0.81794 | 0.09758 | 0.81946 | 0.47735 | 2016-09-19 11:24 | 1.2249 | 1.2258 | 2016-09-19 11:25 | 1.2226 | 1.2235 | 2016-09-19 11:26 | 1.2203 | 1.2212 | 2016-09-19 11:24 | 1.2237 | 1.2246 | 2016-09-19 11:26 | 1.2214 | 1.2223

## param
-0.9 | 1e-05 |  | -1.9 | 0
0 | 1 |  | -1 | 1
0.9 | 1e-05 |  | 1 | 1
|  |  | 1.9 | 0
304.6 | 1e-07 | 310.55 | 1e-06 | 316.7 | 1e-06 | 323.5 | 1e-06 | 323.1 | 1e-06 | 330.5 | 1e-06 | 338 | 1e-06
305.5 | 1000 | 311.45 | 1000 | 317.6 | 1000 | 324.4 | 1000 | 324 | 1000 | 331.4 | 1000 | 338.9 | 1000
306.4 | 1e-07 | 312.35 | 1e-06 | 318.5 | 1e-06 | 326.4 | 1000 | 326 | 1000 | 333.4 | 1000 | 340.9 | 1000
[6 empty] | 327.3 | 1e-06 | 326.9 | 1e-06 | 334.3 | 1e-06 | 341.8 | 1e-06
1.1111
304.6 | 0.001 | 310.55 | 0.001 | 316.7 | 0.001 | 323.5 | 0 | 323.1 | 0.001 | 330.5 | 0.001 | 338 | 1e-07
305.5 | 100 | 311.45 | 100 | 317.6 | 100 | 324.4 | 100 | 324 | 100 | 331.4 | 100 | 338.9 | 100
306.4 | 0.001 | 312.35 | 0.001 | 318.5 | 0.001 | 326.4 | 100 | 326 | 100 | 333.4 | 100 | 340.9 | 100
[6 empty] | 327.3 | 0 | 326.9 | 0.001 | 334.3 | 0.001 | 341.8 | 1e-07
58 | 0
92 | 0
code | number | name | Dobson | earth_radius | LONG | Hemisphere | degs | min | sec | LAT | Hemisphere | deg | min | sec | TimeZone | stn Pressure | elevation | Ozone_Height | LongRadian | LatRadian
IZO | 999 | Izana Observatory,Spain | 83 | 6373593 | 16.4973 | W | 16 | 29 | 25 | 28.2986 | N | 28 | 17 | 55 | 0 | 770 | 2360 | 26000 | 0.28781 | 0.4939
=IF( C13<90,(earth_radius+Ozone_height)/(SQRT(((earth_radius+Ozone_height)^2)-((earth_radius+Elevation)^2)*(SIN(RADIANS(C13)))^2)),-90000)
Public Function AirMass#(CosZenAng As Double) | [14 empty] | p0 | 0.75993
AirMass# = 1# / (CosZenAng + 0.15 / (90 - RadsToDegs#(ArcCos#(CosZenAng)) + 3.885) ^ 1.253)
End Function
Ozone Absorption and Molecular Scattering coefficients for Use
with Dobson ozone Spectrophotometers Beginning 1 January 1992
Wavelength | ALFA |  | BETA | d(BETA)/d(ALFA) | [5 empty] | IUP Corr-factor | B&BP Corr_factor | D&M |  |  | α_Komhyr r adj
(nm) | (atm-cm)-1 |  | (atm)-1 | (atm-cm/atm) |  |  |  |  | A_ pair | 0.9898 | 0.9969 | 0.9889 |  |  | 1.806
[9 empty] | C_pair | 0.9922 | 0.9983 | 0.9845 |  |  | 0.833
305 | 0.489 | [7 empty] | D_pair | 0.9642 | 0.9825 | 0.9562 |  |  | 0.374
325 | 0.375 | [7 empty] | AD_pair | 0.9965 | 1.0006 | 0.9975 |  |  | 1.432
Alfa_A | 1.7876 | Beta_A | 114 | 63 | 1000 |  |  |  | CD_pair | 1.015 | 1.0112 | 1.0076 |  |  | 0.459
|  |  | 0 | 0
308.9 | 0.466 |  | 0 | 0
329.1 | 0.355 |  | 0 | 0
Alfa_B | 1.192 | Beta_B | 111 | 93
|  |  | 0 | 0
311.5 | 0.45 |  | 0 | 0 |  |  | 5
332.4 | 0.341 |  | 0 | 0
Alfa_C | 0.8265 | Beta_C | 109 | 131
|  |  | 0 | 0
317.5 | 0.414 |  | 0 | 0
339.9 | 0.31 |  | 0 | 0
Alfa_D | 0.36061 | Beta_D | 104 | 278
|  |  | 0 | 0
Alfa_AD | 1.427 | Beta_AD | 10 | 7
Alfa_BD | 0.818 | Beta_BD | 7 | 9
Alfa_CD | 0.46588 | Beta_CD | 5 | 11

## Nfunc083
D83RA | D83NA | D83RC | D83NC | D83RD | D83ND | Rby5 | a | c | d | in pretty print
5 | -1.51 | 5 | 2.18 | 5 | 3.26 | 5 | -1.51 | 2.18 | 3.26 | -1.5 | 2.2 | 3.3 | 0.01 | 0.02 | 0.04
10.74 | 4.7 | 10.76 | 8.39 | 10.72 | 9.62 | 15 | 9.2606 | 12.8971 | 14.1824 | 9.2 | 12.9 | 14.2 | -0.060647 | 0.0029402 | 0.017648
16.48 | 10.83 | 16.51 | 14.49 | 16.44 | 15.7 | 25 | 19.7018 | 23.3471 | 24.6275 | 19.7 | 23.3 | 24.6 | -0.001759 | -0.047139 | -0.027526
22.22 | 16.83 | 22.27 | 20.51 | 22.16 | 21.69 | 35 | 29.9065 | 33.5082 | 34.7616 | 29.9 | 33.5 | 34.8 | -0.0065067 | -0.0081693 | 0.038446
27.96 | 22.74 | 28.02 | 26.46 | 27.88 | 27.58 | 45 | 39.8449 | 43.394 | 44.5867 | 39.8 | 43.4 | 44.6 | -0.044871 | 0.0060056 | 0.013291
33.7 | 28.59 | 33.78 | 32.28 | 33.6 | 33.36 | 55 | 49.3404 | 52.8934 | 54.076 | 49.3 | 52.9 | 54.1 | -0.040402 | 0.0065573 | 0.02402
39.87 | 34.8 | 39.99 | 38.5 | 40.01 | 39.73 | 65 | 58.6907 | 62.1867 | 63.3142 | 58.7 | 62.2 | 63.3 | 0.0093361 | 0.013335 | -0.014152
46.12 | 40.93 | 46.25 | 44.6 | 46.26 | 45.8 | 75 | 67.9965 | 71.4015 | 72.4459 | 68 | 71.4 | 72.4 | 0.0034853 | -0.0015358 | -0.045888
52.43 | 46.93 | 52.58 | 50.62 | 52.56 | 51.79 | 85 | 77.243 | 80.6092 | 81.6241 | 77.2 | 80.6 | 81.6 | -0.043043 | -0.0091746 | -0.024141
58.76 | 52.85 | 58.93 | 56.56 | 58.88 | 57.68 | 95 | 86.5778 | 89.8754 | 90.8361 | 86.6 | 89.9 | 90.8 | 0.022224 | 0.024624 | -0.036083
65.01 | 58.7 | 65.22 | 62.39 | 65.17 | 63.47 | 105 | 95.9718 | 99.1641 | 100.0591 | 96 | 99.2 | 100.1 | 0.028156 | 0.03591 | 0.040888
71.68 | 64.91 | 71.96 | 68.6 | 72.14 | 69.83 | 115 | 105.3164 | 108.4291 | 109.276 | 105.3 | 108.4 | 109.3 | -0.016442 | -0.029128 | 0.024023
78.27 | 71.03 | 78.58 | 74.7 | 78.77 | 75.9 | 125 | 114.7278 | 117.7197 | 118.4987 | 114.7 | 117.7 | 118.5 | -0.027761 | -0.019725 | 0.0013422
84.77 | 77.03 | 85.12 | 80.72 | 85.3 | 81.9 | 135 | 124.2697 | 127.0863 | 127.7279 | 124.3 | 127.1 | 127.7 | 0.030284 | 0.013723 | -0.027928
91.14 | 82.95 | 91.54 | 86.67 | 91.69 | 87.78 | 145 | 133.8086 | 136.5037 | 137.0681 | 133.8 | 136.5 | 137.1 | -0.0086133 | -0.0037106 | 0.031879
97.36 | 88.8 | 97.82 | 92.49 | 97.96 | 93.57 | 155 | 143.4602 | 145.9931 | 146.5049 | 143.5 | 146 | 146.5 | 0.03978 | 0.0069477 | -0.0049043
103.97 | 95.01 | 104.5 | 98.7 | 104.86 | 99.93 | 165 | 153.1836 | 155.5627 | 155.9752 | 153.2 | 155.6 | 156 | 0.016436 | 0.037317 | 0.02477
110.54 | 101.13 | 111.08 | 104.8 | 111.45 | 106.01 | 175 | 162.9608 | 165.2327 | 165.5125 | 163 | 165.2 | 165.5 | 0.039212 | -0.032674 | -0.012494
116.94 | 107.14 | 117.58 | 110.82 | 117.96 | 112 | 185 | 172.7866 | 174.9334 | 175.074 | 172.8 | 174.9 | 175.1 | 0.013394 | -0.033433 | 0.026028
123.23 | 113.05 | 123.98 | 116.77 | 124.34 | 117.89 | 195 | 182.6394 | 184.6057 | 184.6243 | 182.6 | 184.6 | 184.6 | -0.039352 | -0.0057191 | -0.024324
129.38 | 118.9 | 130.22 | 122.59 | 130.61 | 123.67 | 205 | 192.4841 | 194.2836 | 194.1583 | 192.5 | 194.3 | 194.2 | 0.015876 | 0.016413 | 0.041746
135.88 | 125.11 | 136.83 | 128.81 | 137.48 | 130.03 | 215 | 202.4109 | 204.0518 | 203.7711 | 202.4 | 204.1 | 203.8 | -0.010885 | 0.048175 | 0.028851
142.31 | 131.24 | 143.3 | 134.9 | 143.98 | 136.11 | 225 | 212.4107 | 213.9099 | 213.4697 | 212.4 | 213.9 | 213.5 | -0.010729 | -0.0099222 | 0.030315
148.58 | 137.24 | 149.67 | 140.92 | 150.34 | 142.1 | 235 | 222.6125 | 223.9066 | 223.3561 | 222.6 | 223.9 | 223.4 | -0.01248 | -0.0066373 | 0.043946
154.69 | 143.16 | 155.92 | 146.87 | 156.57 | 147.99 | 245 | 232.9094 | 234.043 | 233.3305 | 232.9 | 234 | 233.3 | -0.0094379 | -0.04302 | -0.030472
160.72 | 149 | 162.02 | 152.69 | 162.68 | 153.77 | 255 | 243.0375 | 244.1782 | 243.3484 | 243 | 244.2 | 243.3 | -0.037453 | 0.021839 | -0.048358
167.07 | 155.21 | 168.46 | 158.91 | 169.37 | 160.14 | 265 | 253.2588 | 254.2376 | 253.357 | 253.3 | 254.2 | 253.4 | 0.041246 | -0.037564 | 0.043032
173.34 | 161.34 | 174.77 | 165.01 | 175.73 | 166.21 | 275 | 263.5293 | 264.3569 | 263.3423 | 263.5 | 264.4 | 263.3 | -0.029307 | 0.043099 | -0.042337
179.48 | 167.34 | 180.98 | 171.03 | 181.99 | 172.2 | 285 | 273.6505 | 274.5418 | 273.4295 | 273.7 | 274.5 | 273.4 | 0.049496 | -0.041795 | -0.029511
185.48 | 173.26 | 187.11 | 176.98 | 188.16 | 178.09 |  |  |  |  | 284 | 284.8 | 283.6 | -0.080761 | -0.032459 | 0.10211
191.42 | 179.11 | 193.13 | 182.8 | 194.22 | 183.88
197.72 | 185.32 | 199.56 | 189.01 | 200.89 | 190.24
203.94 | 191.44 | 205.85 | 195.11 | 207.25 | 196.31
210.02 | 197.44 | 212.02 | 201.13 | 213.48 | 202.31
215.95 | 203.36 | 218.08 | 207.08 | 219.6 | 208.2
221.81 | 209.21 | 223.98 | 212.9 | 225.52 | 213.98
227.98 | 215.42 | 230.23 | 219.11 | 231.96 | 220.34
233.97 | 221.55 | 236.29 | 225.21 | 238.08 | 226.42
239.77 | 227.55 | 242.22 | 231.23 | 244.08 | 232.41
245.54 | 233.46 | 248.1 | 237.18 | 249.96 | 238.3
251.31 | 239.31 | 253.83 | 243 | 255.73 | 244.08
257.45 | 245.52 | 260.03 | 249.22 | 262.08 | 250.45
263.44 | 251.65 | 266.07 | 255.32 | 268.18 | 256.52
269.25 | 257.65 | 272.01 | 261.34 | 274.17 | 262.51
275.04 | 263.57 | 277.89 | 267.28 | 280.03 | 268.4
280.82 | 269.42 | 283.59 | 273.11 | 285.75 | 274.19
286.94 | 275.63 | 289.71 | 279.32 | 292.01 | 280.55
292.84 | 281.75 | 295.61 | 285.42 | 298.01 | 286.62
298.49 | 287.75

## comp page
Komyhr adjusted |  |  |  |  | IUP multiply |  |  |  |  | B&P B multiply |  |  |  |  | D&M multiply
271.1097 | 281.2914 | 270.7211 | 262.8625 | 267.9448 | 273.241 | 291.7356 | 273.5109 | 258.9778 | 268.8859 | 271.5714 | 286.3017 | 271.5629 | 259.951 | 267.7842 | 275.378 | 294.1764 | 273.7598 | 260.8798 | 268.6164
D083_TO3_C | D083_TO3_D | D083_TO3_A | D083_TO3_CD | D083_TO3_AD | D083_TO3_C | D083_TO3_D | D083_TO3_A | D083_TO3_CD | D083_TO3_AD | D083_TO3_C | D083_TO3_D | D083_TO3_A | D083_TO3_CD | D083_TO3_AD | D083_TO3_C | D083_TO3_D | D083_TO3_A | D083_TO3_CD | D083_TO3_AD
268.0484 | 272.8928 | 268.2957 | 264.3716 | 267.0104 | 270.1557 | 283.0251 | 271.0605 | 260.4646 | 267.9483 | 268.5049 | 277.7535 | 269.13 | 261.4434 | 266.8503 | 272.2686 | 285.393 | 271.3072 | 262.3775 | 267.6796
268.0752 | 273.8684 | 268.4967 | 263.5798 | 267.0232 | 270.1826 | 284.0369 | 271.2636 | 259.6845 | 267.9611 | 268.5317 | 278.7464 | 269.3316 | 260.6604 | 266.8631 | 272.2957 | 286.4133 | 271.5105 | 261.5917 | 267.6924
268.544 | 273.048 | 269.0222 | 265.0634 | 267.9113 | 270.6551 | 283.186 | 271.7945 | 261.1462 | 268.8523 | 269.0013 | 277.9114 | 269.8588 | 262.1275 | 267.7507 | 272.772 | 285.5553 | 272.0419 | 263.0641 | 268.5828
268.1674 | 273.4187 | 268.7526 | 264.049 | 267.4833 | 270.2755 | 283.5706 | 271.5221 | 260.1468 | 268.4228 | 268.624 | 278.2888 | 269.5883 | 261.1244 | 267.3229 | 272.3894 | 285.9431 | 271.7692 | 262.0573 | 268.1537
268.1144 | 273.9415 | 268.8378 | 263.5037 | 267.4615 | 270.2221 | 284.1127 | 271.6082 | 259.6096 | 268.4009 | 268.571 | 278.8209 | 269.6738 | 260.5852 | 267.3011 | 272.3356 | 286.4897 | 271.8554 | 261.5162 | 268.1318
269.3095 | 275.0477 | 269.2623 | 264.7522 | 267.7139 | 271.4266 | 285.26 | 272.0371 | 260.8396 | 268.6542 | 269.7681 | 279.9467 | 270.0996 | 261.8198 | 267.5534 | 273.5495 | 287.6466 | 272.2847 | 262.7553 | 268.3849
269.6271 | 275.5276 | 269.2643 | 264.9219 | 267.596 | 271.7467 | 285.7577 | 272.0391 | 261.0068 | 268.5359 | 270.0862 | 280.4352 | 270.1016 | 261.9877 | 267.4355 | 273.8721 | 288.1485 | 272.2867 | 262.9237 | 268.2667
269.9271 | 277.3121 | 269.8852 | 263.9994 | 267.9171 | 272.0491 | 287.6085 | 272.6664 | 260.0979 | 268.8581 | 270.3867 | 282.2515 | 270.7245 | 261.0753 | 267.7564 | 274.1768 | 290.0147 | 272.9146 | 262.0081 | 268.5885
269.2867 | 277.2629 | 269.7284 | 262.8664 | 267.7355 | 271.4037 | 287.5575 | 272.5079 | 258.9816 | 268.6759 | 269.7453 | 282.2015 | 270.5671 | 259.9549 | 267.575 | 273.5264 | 289.9633 | 272.756 | 260.8837 | 268.4065
270.5513 | 277.1839 | 269.9181 | 265.2165 | 267.9983 | 272.6782 | 287.4755 | 272.6996 | 261.297 | 268.9396 | 271.012 | 282.121 | 270.7574 | 262.279 | 267.8376 | 274.8109 | 289.8807 | 272.9478 | 263.216 | 268.67
269.8651 | 276.3581 | 269.3454 | 264.6362 | 267.4942 | 271.9866 | 286.619 | 272.121 | 260.7253 | 268.4337 | 270.3246 | 281.2805 | 270.1829 | 261.7051 | 267.3338 | 274.1138 | 289.017 | 272.3687 | 262.6401 | 268.1646
270.2533 | 277.9283 | 269.8665 | 264.0545 | 267.7435 | 272.3778 | 288.2475 | 272.6475 | 260.1522 | 268.6839 | 270.7135 | 282.8787 | 270.7057 | 261.1298 | 267.5829 | 274.5081 | 290.6591 | 272.8957 | 262.0628 | 268.4145
270.5982 | 278.1643 | 270.2591 | 264.4825 | 268.1788 | 272.7255 | 288.4923 | 273.0441 | 260.5739 | 269.1207 | 271.059 | 283.1188 | 271.0995 | 261.5531 | 268.018 | 274.8586 | 290.9059 | 273.2926 | 262.4876 | 268.8509
270.4784 | 278.158 | 270.1566 | 264.2652 | 268.0528 | 272.6048 | 288.4858 | 272.9406 | 260.3598 | 268.9943 | 270.939 | 283.1125 | 270.9967 | 261.3382 | 267.892 | 274.7369 | 290.8994 | 273.189 | 262.2719 | 268.7246
270.8788 | 279.1584 | 270.3582 | 264.1724 | 268.0471 | 273.0083 | 289.5233 | 273.1442 | 260.2683 | 268.9885 | 271.3401 | 284.1307 | 271.1989 | 261.2464 | 267.8863 | 275.1436 | 291.9456 | 273.3928 | 262.1798 | 268.7189
269.9798 | 277.8908 | 270.153 | 263.5699 | 268.1206 | 272.1022 | 288.2086 | 272.937 | 259.6747 | 269.0623 | 270.4396 | 282.8405 | 270.9931 | 260.6506 | 267.9598 | 274.2304 | 290.6199 | 273.1854 | 261.5818 | 268.7926
269.7764 | 279.0684 | 270.0379 | 262.2378 | 267.669 | 271.8972 | 289.43 | 272.8207 | 258.3623 | 268.6091 | 270.2358 | 284.0391 | 270.8777 | 259.3332 | 267.5084 | 274.0237 | 291.8515 | 273.069 | 260.2598 | 268.3398
270.4101 | 278.052 | 270.0503 | 264.2131 | 267.951 | 272.5359 | 288.3759 | 272.8332 | 260.3085 | 268.8921 | 270.8706 | 283.0046 | 270.8901 | 261.2867 | 267.7903 | 274.6674 | 290.7885 | 273.0816 | 262.2202 | 268.6226
270.9323 | 278.5825 | 270.5511 | 264.7259 | 268.4448 | 273.0622 | 288.9261 | 273.3391 | 260.8137 | 269.3876 | 271.3937 | 283.5445 | 271.3924 | 261.7938 | 268.2838 | 275.1979 | 291.3433 | 273.5879 | 262.7292 | 269.1176
270.8273 | 278.2479 | 270.033 | 264.8056 | 267.8796 | 272.9563 | 288.579 | 272.8158 | 260.8923 | 268.8205 | 271.2885 | 283.204 | 270.8727 | 261.8727 | 267.7189 | 275.0912 | 290.9934 | 273.064 | 262.8083 | 268.551
270.0912 | 281.1438 | 270.6597 | 261.1081 | 267.9142 | 272.2144 | 291.5825 | 273.4489 | 257.2493 | 268.8552 | 270.5511 | 286.1515 | 271.5013 | 258.216 | 267.7536 | 274.3435 | 294.022 | 273.6977 | 259.1386 | 268.5857
271.6366 | 280.6145 | 270.3541 | 264.3421 | 267.6676 | 273.772 | 291.0335 | 273.1401 | 260.4356 | 268.6078 | 272.0991 | 285.6127 | 271.1948 | 261.4143 | 267.5071 | 275.9132 | 293.4684 | 273.3887 | 262.3483 | 268.3385
271.52 | 282.1865 | 269.8874 | 262.8481 | 266.669 | 273.6546 | 292.6639 | 272.6686 | 258.9637 | 267.6056 | 271.9824 | 287.2127 | 270.7266 | 259.9368 | 266.5091 | 275.7949 | 295.1124 | 272.9167 | 260.8655 | 267.3373
270.2608 | 280.2025 | 270.3976 | 262.1779 | 267.8311 | 272.3854 | 290.6062 | 273.1841 | 258.3034 | 268.7718 | 270.721 | 285.1934 | 271.2384 | 259.2741 | 267.6705 | 274.5158 | 293.0375 | 273.4327 | 260.2004 | 268.5023
270.9363 | 281.9798 | 269.6535 | 261.9543 | 266.4289 | 273.0662 | 292.4495 | 272.4323 | 258.083 | 267.3647 | 271.3977 | 287.0024 | 270.492 | 259.0529 | 266.2692 | 275.2019 | 294.8963 | 272.6803 | 259.9784 | 267.0967
271.1778 | 283.2195 | 270.3161 | 261.3813 | 266.9412 | 273.3097 | 293.7352 | 273.1017 | 257.5185 | 267.8787 | 271.6396 | 288.2641 | 271.1567 | 258.4863 | 266.7811 | 275.4473 | 296.1928 | 273.3503 | 259.4098 | 267.6102
272.2892 | 283.9244 | 270.0722 | 262.8228 | 266.4498 | 274.4298 | 294.4663 | 272.8553 | 258.9387 | 267.3856 | 272.7529 | 288.9815 | 270.912 | 259.9118 | 266.29 | 276.5761 | 296.9299 | 273.1036 | 260.8404 | 267.1176
270.3554 | 284.0974 | 270.7206 | 259.1714 | 267.2227 | 272.4808 | 294.6457 | 273.5104 | 255.3412 | 268.1612 | 270.8158 | 289.1577 | 271.5624 | 256.3008 | 267.0624 | 274.6119 | 297.1109 | 273.7593 | 257.2165 | 267.8924
272.692 | 285.3782 | 270.7375 | 262.3674 | 266.9098 | 274.8358 | 295.9741 | 273.5275 | 258.49 | 267.8473 | 273.1564 | 290.4613 | 271.5794 | 259.4614 | 266.7498 | 276.9853 | 298.4503 | 273.7764 | 260.3884 | 267.5788
273.0186 | 281.244 | 270.6966 | 266.3236 | 267.9396 | 275.1649 | 291.6863 | 273.4862 | 262.3878 | 268.8807 | 273.4835 | 286.2534 | 271.5384 | 263.3738 | 267.779 | 277.317 | 294.1267 | 273.7351 | 264.3148 | 268.6112
271.8231 | 280.4626 | 270.7157 | 264.7902 | 268.168 | 273.96 | 290.876 | 273.5055 | 260.8771 | 269.1098 | 272.286 | 285.4582 | 271.5576 | 261.8574 | 268.0071 | 276.1027 | 293.3096 | 273.7544 | 262.793 | 268.8401
272.7724 | 279.3171 | 270.5391 | 267.4459 | 268.2444 | 274.9167 | 289.6879 | 273.327 | 263.4935 | 269.1866 | 273.2369 | 284.2922 | 271.3803 | 264.4837 | 268.0836 | 277.0669 | 292.1116 | 273.5758 | 265.4286 | 268.9167
271.8353 | 281.7526 | 270.9083 | 263.7604 | 268.0741 | 273.9723 | 292.2139 | 273.7 | 259.8625 | 269.0157 | 272.2982 | 286.7711 | 271.7507 | 260.839 | 267.9134 | 276.1151 | 294.6587 | 273.9491 | 261.771 | 268.746
273.1538 | 283.9393 | 271.8479 | 264.3711 | 268.6882 | 275.3011 | 294.4818 | 274.6493 | 260.4641 | 269.6319 | 273.6189 | 288.9967 | 272.6933 | 261.4429 | 268.5271 | 277.4543 | 296.9455 | 274.8993 | 262.377 | 269.3616
273.4208 | 283.8577 | 271.8652 | 264.9219 | 268.7314 | 275.5702 | 294.3971 | 274.6668 | 261.0068 | 269.6752 | 273.8864 | 288.9136 | 272.7106 | 261.9876 | 268.5702 | 277.7255 | 296.8601 | 274.9167 | 262.9237 | 269.4049
271.7369 | 283.4802 | 272.4706 | 262.1732 | 269.5936 | 273.8731 | 294.0056 | 275.2785 | 258.2987 | 270.5405 | 272.1996 | 288.5295 | 273.3179 | 259.2694 | 269.432 | 276.0151 | 296.4654 | 275.529 | 260.1957 | 270.2693
275.0511 | 295.2295 | 274.4647 | 258.6141 | 269.04 | 277.2134 | 306.1911 | 277.2931 | 254.7923 | 269.9849 | 275.5195 | 300.488 | 275.3181 | 255.7498 | 268.8786 | 279.3816 | 308.7529 | 277.5454 | 256.6635 | 269.7142
276.4619 | 296.5086 | 274.372 | 260.1318 | 268.5891 | 278.6352 | 307.5177 | 277.1994 | 256.2875 | 269.5324 | 276.9326 | 301.7899 | 275.2252 | 257.2506 | 268.428 | 280.8145 | 310.0906 | 277.4517 | 258.1698 | 269.2622
274.1104 | 295.4313 | 274.9049 | 256.7419 | 269.5426 | 276.2652 | 306.4004 | 277.7378 | 252.9477 | 270.4893 | 274.5772 | 300.6934 | 275.7597 | 253.8982 | 269.381 | 278.426 | 308.9639 | 277.9906 | 254.8054 | 270.2182
275.1392 | 296.2067 | 274.4954 | 257.9771 | 268.8237 | 277.3022 | 307.2046 | 277.3241 | 254.1646 | 269.7679 | 275.6078 | 301.4826 | 275.3489 | 255.1197 | 268.6625 | 279.471 | 309.7748 | 277.5765 | 256.0312 | 269.4974
268.6374 | 290.265 | 274.328 | 251.0186 | 270.1646 | 270.7493 | 301.0423 | 277.155 | 247.3089 | 271.1135 | 269.0949 | 295.4351 | 275.1811 | 248.2383 | 270.0026 | 272.8669 | 303.561 | 277.4073 | 249.1252 | 270.8417
274.8358 | 292.6876 | 273.6033 | 260.2934 | 268.6179 | 276.9964 | 303.5549 | 276.4228 | 256.4467 | 269.5614 | 275.3039 | 297.9009 | 274.4541 | 257.4104 | 268.4569 | 279.1629 | 306.0945 | 276.6744 | 258.3301 | 269.2912
274.4055 | 288.3133 | 273.5359 | 263.0764 | 269.6754 | 276.5627 | 299.0182 | 276.3547 | 259.1886 | 270.6226 | 274.8728 | 293.4487 | 274.3865 | 260.1626 | 269.5137 | 278.7257 | 301.5199 | 276.6062 | 261.0921 | 270.3513
268.0899 | 285.8528 | 274.1475 | 253.6195 | 271.0893 | 270.1975 | 296.4663 | 276.9726 | 249.8714 | 272.0415 | 268.5465 | 290.9444 | 275 | 250.8104 | 270.9268 | 272.3108 | 298.9467 | 277.2247 | 251.7065 | 271.7688
272.9972 | 287.6148 | 274.2133 | 261.0894 | 270.7122 | 275.1433 | 298.2937 | 277.0391 | 257.231 | 271.663 | 273.4621 | 292.7377 | 275.066 | 258.1976 | 270.5499 | 277.2953 | 300.7894 | 277.2912 | 259.1201 | 271.3907
272.991 | 289.192 | 273.723 | 259.7929 | 269.6821 | 275.1371 | 299.9294 | 276.5438 | 255.9536 | 270.6293 | 272.5269 | 284.1311 | 272.8745 | 262.7025 | 269.8439 | 277.289 | 302.4388 | 276.7955 | 257.8333 | 270.358