## Slide 1
Instruments/Measuring Aerosols
Elisabeth Andrews
CATCOS Aerosol Workshop, June 2016

## Slide 2
Aerosol Property | Instruments
Mass | Filters, TEOM…
Composition | Filters, AMS…
Number | CPC
Size | SMPS, OPC…
Scattering | Nephelometer
Absorption | PSAP, Aethalometer…
Making aerosol measurements
Different instruments used to measure different particle properties.
May need multiple instruments to answer science questions
Wide range in complexity of aerosol instruments
Instruments measure on different time scales (<seconds to >days)
Note: there are other aerosol properties that can also be measured (e.g., shape, mixing state, hygroscopicity, …) – those are not covered here!

## Slide 3
In-situ
Offline – time integrated sample over extended time (e.g., PM10)
Online – ~continuous, ~real time measurement (e.g., aethalometer)
Remote
Column – integrated through the atmosphere (e.g., AOD)
Profile – function of altitude (e.g., Lidar)
Categories of aerosol measurements
In-situ instrument
Remote instrument
Particles collected/sensed inside instrument
Effect of particles sensed (typically sun or laser based)
Notes: Could also split into active and passive. Passive is primarily for gases, typically uses adsorption; active uses pumps to pull air through something

## Slide 4
Categories of in-situ aerosol measurements
AMS, XRF, PILS, EC/OC
Particles collected on filter
Filters extracted
Extract sent through chemical analyzer
Bulk analysis
Integrated over time
Particles directly sampled by chemical analyzer
Can be bulk or single particle
Higher time resolution than filters
From Jimenez AMS webpages
Offline Analysis
Online Analysis

## Slide 5
Volume of air = Qsample
tsample
Aerosol Mass – offline method
1. In the lab
Clean filter is weighed (Mclean)
2. At the site
Specially shaped inlet – size cut Particles collected onto a filter
Constant flow rate (Qsample)
Sample time recorded (tsample)
3. In the lab
RH equilibration of filter
Sampled filter is weighed (Msample)
PM10 = Particle mass
Volume
Particle mass = (Mclean-Msample)

## Slide 6
From Hadjimmitsis et al. 2013
Aerosol Mass – Example offline data
Filter measurements typically collected over hours to days.
Plot shows data for ~6 months
Samples collected every 5-10 days

## Slide 7
Tapered-Element Oscillating Microbalance (TEOM)
Continuous real time measurement
Particles collect on vibrating substrate
Oscillation changes with mass
Aerosol Mass – Online methods
Inlet
Vibrating
Substrate
Electronics
Inlet
Filter tape
Beta radiation source and detector
Beta Attenuation Monitor
(BAM)
Hourly time measurement
Particles collect on filter tape
Beta ray penetration changes with mass

## Slide 8
Aerosol Mass – Example online data
Advantage of high frequency measurements:
captures variability and exposure on shorter time scales
Disadvantage of high frequency measurements:
more complex instrumentation – more scientist time, more cost
PM2.5 concentration (mg/m3)
Plot shows data for ~2 days
Samples collected every hour

## Slide 9
Aerosol chemistry – offline
Chemical analysis of filter samples can require multiple steps after sample collection.
Filter with aerosol particles collected on it.
Extract particles into liquid
Filter to remove non-soluble particles
Ion chromatograph
Inject soluble fraction into chemical analyzer

## Slide 10
Examples of chemical filter analysis methods:
Inorganic ions (ion chromatograph)
EC/OC (Thermogram)
Elements (PIXE, XRF)
---Temperature
---Carbon as CO2
Aerosol chemistry – offline

## Slide 11
Aerosol chemistry – size-segregated (offline)
From http://elte.prompt.hu
Cascade impactor
Multiple levels
Each levels collects particles of different size
Substrates (filters) analyzed offline
Bulk measurement
May require long sampling times
Can also use for mass
Smallest size cut typically <0.4 mm

## Slide 12
Aerosol chemistry – size segregated (offline)
From Sandrini et al., 2016
Diameter
Mass Concentration
Variations in chemistry and size distribution with time.
Smallest size typically <0.4 mm

## Slide 13
Mass spectrometry (MS) is an analytical technique that ionizes chemical species and sorts the ions based on their mass-to-charge ratio.
From Max Planck Inst
From WIkipedia
Aerosol Chemistry - Online
Aerosol Mass Spectrometry

## Slide 14
Aerosol Chemistry – Example online data
Time series of chemical concentration
Bulk chemical composition
Size distribution of chemical composition

## Slide 15
Aerosol Number – Online
Photodetector
Laser
TSI Inc
Particles exposed to supersaturated vapor
Particles grow to size that can be optically detected
Minimum detection size related to instrument temperatures/supersaturations
Minimum size ~0.01 mm
Condensation Particle Counter
Notes: Size ran

## Slide 16
TSI 3010 dp> 0.01 mm
TSI 3776 dp> 0.004 mm
Day of Year, 2016
Number concentration (cm-3)
Aerosol Number – Example data
Significant difference between two particle counters due to new particles formed by nucleation “secondary aerosol”
Gas phase
molecules
Molecular
Cluster
Detectable
Particle
Nucleation

## Slide 17
Aerosol Size Distribution – Optical Method
Light
source
Minimum size ~0.1 mm
Particles scatter light
Amount of scattering related to size
Calibrated with particles of known size
Optical Particle Counter

## Slide 18
Aerosol Size Distribution – Electrical Mobility Method
Differential Mobility Analyzer (DMA)
particles charged
amount of charge depends on size
voltage determines size exiting
particles are counted
Size range ~0.01-1 mm
Particle counter
High voltage rod
DMAs with changing voltages are called:
SMPS (scanning mobility particle sizer)
or
SEMS (scanning electrical mobility sizer)

## Slide 19
Aerosol Size Distribution – Example data
From J. Hand
From Covert et al 1996
‘banana’ curve used to identify nucleation events
Merged volume size distribution data from OPC and SMPS
DMA
OPC
Size distributions change with time

## Slide 20
Aerosol Optical Properties - Scattering
Nephelometer
particles scatter light in sensing volume
photomultiplier tubes detect scattered light
Common options:
Multiple wavelength
Angular dependence of scattering
Calibrated with gas like CO2

## Slide 21
AVHRR satellite image
scattering (Mm-1)
Aerosol Optical Properties – Example scattering data
Surface and airborne nephelometer measurements show passage of smoke plume.

## Slide 22
Aerosol Optical Properties - Absorption
Filter-based absorption measurements:
PSAP – Particle Soot Absorption Photometer
TAP – Tri-color Absorption Photometer
Aethalometer
MAAP – Multi-Angle Absorption Photometer
Light source
MAAP
TAP filters
These are on-line measurements

## Slide 23
Aerosol Optical Properties - Absorption
Light source
Particle absorbs light
Heating
Pressure (sound) wave
Acoustical detection
The heated air expands, resulting in sound wave
From Arnott et al 2005
1
2
3
4
Photo-acoustic Spectrometer
advantage: no scattering or filter effects to interfere with measurements
disadvantage: can be sensitive to pressure changes, complicated!

## Slide 24
From Arnott 2005
Aerosol Optical Properties – Example absorption data
Filter-based PSAP and photo-acoustic spectrometer track each other.
PSAP > photo-acoustic because PSAP has scattering artifact

## Slide 25
Aerosol Optical Properties – Pha Din
Measurements started, 2014
Aerosol light absorption
Aerosol light scattering
Aethalometer
Nephelometer

## Slide 26
Remote sensing – Aerosol Optical Depth
Sunphotometer
The Sun is the light source
The atmosphere is the ‘sensing volume’
Sun photometer points directly at sun
Light detector measures how much solar radiation comes through atmosphere
As the aerosol increases, more light is scattered out of the direct solar beam  sun photometer measurement decreases.
Daytime measurement only

## Slide 27
Remote sensing – Aerosol extinction profile
Pulsed laser
Particles scatter light
Backscattered light
Lidar
Pulsed laser is the light source
The atmosphere is the ‘sensing volume
Telescope detects backscattered light
Because laser is pulsed can use time of backscatter returns to determine height

## Slide 28
Remote sensing – Aerosol extinction profile
Lidar
Pulsed laser is the light source
The atmosphere is the ‘sensing volume
Telescope detects backscattered light
Because laser is pulsed can use time of backscatter returns to determine height

## Slide 29
Surface-based lidar measures as function of time
Airborne/satellite lidar measures as function of distance
From MPLNet lidar
From NASA satellite lidar
Remote sensing – Lidar example data

## Slide 30
Zieger et al 2014
Source region and water uptake
Scattering
Chemistry and size
Aerosol Measurements – Putting them all together!

## Slide 31
Many methods to measure different characteristics of aerosol particles
Multiple methods to measure the same aerosol characteristic
Choose methodology based on:
science
cost
time

## Slide 32
Unused slides

## Slide 33
From Jimenez AMS webpages
Aerosol chemistry – bulk vs single particle

## Slide 34
Sanchez de la Campa et al., 2011
Hourly PIXE and 12 hourly filters

## Slide 35
Extra stuff

## Slide 37
Provides size- and RH-controlled measurements of aerosol
  light scattering
  light absorption
  number concentration
Common, core design at NOAA-federated stations
Optional components: CCN, humidified nephelometry (f(RH)), SMPS, aerosol composition
NOAA Aerosol Rack at Mauna Loa, Hawaii