|  | Droplet Measurement Technologies Inc. 2400 Trade Centre Avenue, Longmont, CO, 80503, USA Ph: 303-440-5576, Fax: 303-440-1965 www.dropletmeasurement.com | CCN CAL Data |
| --- | --- | --- |
| Customer: NOAA/U.C. San Diego | Project: CCN-100 | Work Order No.: S17391A |
| Serial No. CCN: 0604-001 | Technician: dsj | Date: 4/26/2022 |
| Labels: | Serial Tag: | Software Rev Level: 7.0.1.100 |
| OPC SN: H347 | Win 10 OS | CDPE Firmware Rev: N/A |
| Computer Speed: 1.9GHZ | Hard drive: 256gb |  |
| Leak Rate: < 1 mb per sec. | Power Supplies: | 0Ω Between All GNDs: |
| Set 2.17V: | Check +7V: 7.66V | Check Fuse Values: |
| Fans: | LEDS: | Diamond card Calibration: |
| Power Resistors: 3Ω 50W | Power switch: 20A |  |
| D.C. Voltage Set: | D.C. Voltage Read: | Temperatures Calibrated: |
| Solenoid Pumps: 60ML 20ML |  |  |
| Valve Min: 0.1 | Valve max: 4.9 |  |
| Sample Flow Slope: 80.0313 | Sheath Flow Slope: 1121.82 | Pressure Slope: 275.829 |
| Sample Flow Int.: -158.694 | Sheath Flow Int.: -2239.76 | Pressure Int.: 112.805 |
| Current T1: 3.17A | Current T2: 3.17A | Current T3: 3.19A |
| T1 Gain: 1.999 | T2 Gain: 1.993 | T3 Gain: 2.007 |
| T1 Offset: 0 | T2 Offset: 0 | T3 Offset: 0 |
| TG %: 7 | DMA Calibration Slope: 13.05 | DMA Calibration INT: 1.73 |
| P.I.D, Values: P:1 I:1 D:0 |  |  |
| Laser Power: 22mW | Laser Current: 90mA | Current limit: 3.43V |
| Current Alarm: 105 mA | Baseline 0 level: 2.66V | Baseline calibrated: 2.76V |
| OPC 1st Stage: 0.14V | CDPE +-5V: 5.02V | 1st stage Rail: 4.95V |
| Chiller ON: 22K | Chiller OFF: 26K | OPC Air Pump: |
| Water Pumps: | 60 Min Zero Count: | 48 Hour burn in: |
| Cold Turn on: |  |  |
| USB Port: | Data Out Port: | Ethernet Port: |
| Touch Screen Cal.: | Computer Directory: | Computer Start Menu: |
| Keyboard Verification: | 28V PS verification: |  |
| Bottle Caps Secured: | Column Dried: | Final Visual Inspection: |
| Wires Dressed neatly: | USB Stick: | Torque seal Pots: |
| Accessories: Monitor, Keyboard, +28V Power Supply Comments: Calibrated CCN, calibration data packet included with CCN. Items Replaced: New Column, Heatsinks, Column TEC’s, customer supplied TEC’s Replaced all wiring and tubing WIN-10 Computer, New control, Power boards, I.O. Panel New TEC controller 3 pack Pressure sensor assembly, original one was missing New manifold with ribbon thermistor New 60ml solenoid supply, OPC, Chiller, Column drain pumps New air pump, one supplied from customer was nonoperational New power switch, customer supplied New Bottle Holder, Bottle Hold down Replaced all column fans Replaced side panel fans Rebuilt OPC including new optics, New Laser Replaced OPC isolator, cone Replaced Nafion Replaced Sheath Filter SERV-0018 Warranty of repairs: Warranty of 6 months on any repair work performed or parts replaced (except consumable items). Customer pays shipping to DMT. DMT will pay return shipping via standard ground. |

Comments (continued):