## Sheet1
CCN SN: | 1 | DATE: | 2022-04-21 | Tech: | pk
Sample flow | [6 empty] | NOTE: take 5 points only in range of 20-75, one must be 45 +-2, RSQ must be >.99
|  |  |  | slope | 80.0313
|  |  |  | intercept | -158.6944
Volts |  | Flow |  | rsq | 0.99942
2.24 |  | 20.17 |  |  |  |  | 20
2.39 |  | 32.44 |  |  |  |  | 32
2.54 |  | 45.37 |  |  |  |  | 45
2.73 |  | 60.04 |  |  |  |  | 60
2.93 |  | 75.31 |  |  |  |  | 75
Pressure calibration |  |  |  | slope | 275.8295 |  | Note: take only 5 points in range of 400mb to 1000mb, one will be around 840, one will be around 1000, other 3 will be less than 840, RSQ.must be >.99
|  |  |  | intercept | 112.8054
Volts |  | press mbar |  | rsq | 0.99848
1.05 |  | 402 |  |  |  |  | 400
1.44 |  | 511 |  |  |  |  | 510
1.84 |  | 620 |  |  |  |  | 620
2.24 |  | 730 |  |  |  |  | 730
2.62 |  | 835.9 |  |  |  |  | 840
3.13 |  | 1005 |  |  |  |  | 1000
Sheath flow |  |  |  |  | slope | 1121.8169 | NOTE: take 5 points only in range 200-750, one must be 450 +-10, RSQ.must be >.99
[5 empty] | intercept | -2239.7651
Volts | sample | sheath | total |  | rsq | 0.9974
2.17 | 13.14 | 186.66 | 199.8 |  |  |  | 200
2.26 | 19.31 | 306.29 | 325.6 |  |  |  | 325
2.37 | 26.65 | 423.85 | 450.5 |  |  |  | 450
2.51 | 36.95 | 562.05 | 599 |  |  |  | 600
2.62 | 44.4 | 705.6 | 750 |  |  |  | 750
|  |  |  | Flow Meter | CCN Display | % | Set total flow to 500, flow ratio to 10 and measure flows, sample must be +-2.5%, Total must be +-5%
|  |  | Sample | 44.47 | 44.18 | 0.65213
|  |  | Total | 501 | 499.42 | 0.31537
Max Flow ratio Must be 10.6-10.8, adjust restriction if needed
Sample flow rail >150
5/6/2015 dsj

## Flow calibration instructions
Sample flow calibration
1.close sheath flow metering valve
2. enter 1 for slope and 0 for intercept into the sample flow coefficients
3. adjust the flow manually with the valve voltage to get sample flows in the range of 20-100 VCCM. Measure the flow with a flow meter connected to the inlet. Record the measured flows and the displayed voltage (in the flow box) into the spread sheet.
4. Verify RSQ is >.99
5 enter the new slope and intercept values into the sample flow coefficients.
6. Adjust valve voltage to get sample flow reading of 45, verify unit reading agrees with external flow meter before proceeding with sheath flow calibration.
Sheath flow calibration
1. Open sheath flow valve
2. enter 1 for slope and 0 for intercept into the sheath flow coefficients
3. adjust the flow manually with the valve voltage to get a total flows in the range of 200-750 VCCM. Measure the flow with a flow meter connected to the inlet. Record the measured total flows and the displayed voltage (in the flow box) and the displayed sample flow into the spread sheet.
4. Verify RSQ is >.99
5 enter the new slope and intercept values into the sheath flow coefficients.
6. Set total flow to 500 and flow ratio to 10, verify measured total flow agrees with displayed total flow.