## Slide 1: Update and expansion of the WMO X2004 CH4 mole fraction scale
Ed Dlugokencky1, Andrew Crotwell1,2, Patricia Lang1,
Kirk Thoning1, and Brad Hall1
1NOAA Earth System Research Laboratory,
Global Monitoring Division, Boulder, CO, USA
2Also at University of Colorado, CIRES, Boulder, CO USA
(ed.dlugokencky@noaa.gov)

## Slide 2: Outline
Reasons for changes
  Ensure continental measurements are calibrated
  Account for analytical non-linearity
New gravimetrically prepared 1° standards
  6 standards in range 2200 to 5900 ppb
  Combined with previous set (300 to 2600 ppb)
New 2° standards (390 to 5000 ppb)
  Calibrated with gravimetrically-prepared 1°s
  Propagate scale with response curve
How do you update your assigned values?

## Slide 3
Source: Dlugokencky et al., JGR, doi:10.1029/2005JD006035, 2005.
X2004 CH4 Primary Standards

## Slide 5
5
Status mid-June, 2013
Grav value – measured value
(Based on linear response to original 300 - 5760 ppb 1° standards)

## Slide 6
Normalized Response
Tested:
- Different transfer methods
- Inclusion of 10 and 20 ppm standards
- Original 5.8 ppm standard was not consistent with others
5.8 ppm

## Slide 7
Cylinder ID | Year Prepared | Assigned CH4 (ppb) | Uncertainty* (ppb)
FA01024 | 1995 | 302.3 | 2.0
FA01028 | 1994 | 478.3 | 2.1
FA01007 | 1995 | 1021.2 | 2.5
FF37050 | 1993 | 1457.6 | 2.5
FF37061 | 1993 | 1581.0 | 2.3
FF37040 | 1992 | 1690.1 | 2.3
FF37043 | 1993 | 1736.3 | 3.6
FF37052 | 1992 | 1748.1 | 2.3
FF37056 | 1993 | 1751.6 | 2.3
FF37045 | 1993 | 1780.6 | 2.2
FF37036 | 1991 | 1837.1 | 2.4
FA01017 | 1995 | 1896.3 | 2.4
FF37057 | 1992 | 1918.4 | 2.4
FF37024 | 1992 | 2010.5 | 2.3
FB03575 | 2013 | 2227.1 | 1.8
FF39454 | 1995 | 2322.4 | 2.6
FF39474 | 1995 | 2614.5 | 2.7
FF32820 | 2013 | 3140.4 | 2.0
FA02368 | 2013 | 3727.1 | 2.2
FF32783 | 2013 | 4334.0 | 2.5
FF32800 | 2013 | 5018.5 | 2.7
FB03568 | 2013 | 5893.6 | 3.6
22 Primary Standards (X2004A)

## Slide 8
Fit to primary standards + residuals

## Slide 9
Cylinder ID | Year Prepared | Assigned CH4 (ppb) | Uncertainty (ppb) (68% CI)
CB09868 | 2013 | 389.3 | 0.3
CB09893 | 2013 | 849.5 | 0.3
CB09867 | 2013 | 1236.6 | 0.3
CB09898 | 2013 | 1626.7 | 0.3
CB09871 | 2013 | 1797.1 | 0.4
CB09891 | 2013 | 1952.4 | 0.5
CB09804 | 2013 | 2090.6 | 0.3
CB09888 | 2013 | 2433.5 | 0.4
CB09899 | 2013 | 2865.3 | 0.6
CB09900 | 2013 | 3317.5 | 0.4
CB09895 | 2013 | 3762.8 | 0.5
CB09889 | 2013 | 4155.1 | 0.8
CB08856 | 2013 | 4568.9 | 0.8
CB10121 | 2013 | 4998.0 | 1.0
14 Secondary Standards

## Slide 10
Fit to secondary stnds + residuals

## Slide 11: Recalibration of historical secondary standards
All but 2 (of 21) archived in lab
Measured 4X each, 3 GC and 1 CRDS
Mean difference (2004A - 2004): -0.3 ± 0.3 ppb
All calibrations since 17 July, 2013 use RC
  Non-linearity greater since October, 2013
Difference in tertiary standards (300-2600 ppb)
  1983 to September, 2013: -0.4 ± 0.3 ppb
  October, 2013 to present: -0.4 ± 1.9 ppb

## Slide 12
Converting from X2004 to X2004A
http://www.esrl.noaa.gov/gmd/ccl/refgas.html
12

## Slide 13: CH4: Summary and Conclusions
Transitioned from X2004 to X2004A scales
  Expand scale with new gravs at 2200 to 5900 ppb
New secondary standards (390 to 5000 ppb)
  Assigned values based on 6 response curves form primary standards
  Non-linear response: modeled as power function
All historical secondary standards reassigned
  Changes propagated to all calibrations from 1983 to present

## Slide 15
Recalibration: 6 mo or greater
Mean = 0.7 ppb
n = 963
91% are 0 to 1.0 ppb
Recalibration: 3 yr or greater
Mean = 0.6 ppb
n = 499
88% are 0 to 1.0 ppb
CH4 CCL Goal:
Transfer scale to <±1 ppb
All standards

## Slide 16
16
Recalibration: 6 mo or greater
Mean = 0.3 ppb
n = 757
98% are 0 to 1.0 ppb
Recalibration: 3 yr or greater
Mean = 0.3 ppb
n = 392
97% are 0 to 1.0 ppb
CH4 CCL Goal:
Transfer scale to <±1 ppb
Ambient (1600-1900 ppb)

## Slide 18: CH4 Isotope Effects
ΔX ≈ – (X/1000)•13Rstd• Δ(δ13R) – (X/1000)•4•2Rstd• Δ(δD)
≈ – 0.020 ppb ‰-1 for Δ(δ13R), and – 0.00011 ppb ‰-1 for Δ(δD) at ~2 ppm CH4

## Slide 19
Long-term target tanks – 2011
reprocessed using response curve based on primaries