## Slide 1
30 minute flight time with 5 lb payload
Up to 13-pound payload
Enough payload to collect whole air samples for offline VOC analysis
13 minute flight time with 5 lb payload
Easy-to-handle wheeled case
Quick setup/tear down
DJI Matrice 600 Pro
DJI Matrice 210
Unmanned Drones
All of our drones are equipped with winds, temperature, relative humidity, pressure, and methane
Best for near-field quantification, and surface to 400 ft agl flight altitudes
Real-time data display and upload, including images and video

## Slide 2: Measuring the Wind
Scientific Principle
Compare the drone’s actual acceleration with what we predict using the balance of forces
Using the drag force law, calculate the “missing” force and convert that to a wind direction and speed.
Initial Objective - Measure the wind without the addition of ANY hardware
Comparison of winds simultaneously measured by drone and 2-D sonic anemometer

## Slide 3: Measuring Methane Aboard the M210
With custom lightweight case and removal of batteries, we were able to get the Aeris under 4 pounds.
Sensor Requirements
Low Power (~ 15 watts)
Maximum weight 4 pounds
Aeris Pico Specifications
1 ppb CH4 precision
Low power (15 W)
Weight 6 pounds

## Slide 4: High-Level System Architecture
Onboard methane sensor continuously collects and reports methane mixing ratio
iPad app receives raw data from onboard computer and video feed. Data are stored in files and sent to computer (Swift)
Onboard computer collects methane and UAV telemetry and transmits through remote to iPad (C++)
Python app on computer calculates emission rates and generates shape files (Python)

## Slide 5: Test #1: Controlled Release November 5, 2018, Boulder Municipal Airport
Question: Can we see a small leak from a safe distance?
Test Setup:
Obtained air tank with 5% methane (below LEL)
Used flow meter to maintain constant flow rate ~20 liters per minute.
Release rate = 2 scf/hour

## Slide 6: First Test – Operator View
November 5, 2018. We released 20 liters/per minute of 5% methane mixture. That equates to 42 grams/hour or 2 scf/hour.

## Slide 7: First Test – Emissions Analysis
Average flux from 6 laps = 7.8 g/m/hr. Maximum Height = 4.7 m.
Emission = 7.8 g/m/hr * 4.7 m = 37 grams / hour = 1.8 scf/hr